## Radio
Mooring Name | Deployment | Location | Inventory info
|  |  | tBeg | tEnd | OASIS Can | Radio | Radio IP | Comments
OA1 | 201203 | Hopkins
| 201306 | Hopkins
| 201401 | Hopkins | 2014-01-23T23:46:00Z | 2016-07-20T16:20:00Z | "OA2" | "OA1" |  | Can OA2 did NOT have a CO2-standard bulkhead connector for this deployment.
| 201607 | Hopkins | 2016-07-20T19:15:00Z | 2017-02-14T16:00:00Z | "OA5" | "OA3" |  | Can OA5 has a CO2-standard connector.
| 201607 | Hopkins | 2017-02-14T16:01:00Z | 2018-10-08T19:00:00Z |  | "OA RADIO 2" | 134.89.45.24:20000 | Recovered "OA3" to install it in new-style housing. Radio had 192.168.102.1 written on it?
| 201810 | Hopkins | 2018-11-16T10:30:00Z | 2020-10-20T19:00:00Z | "OA5" | "OA RADIO 2" | 134.89.45.24:20000
| 202010 | Hopkins | 2020-10-20T23:25:00Z |  | "OA6" | "OA RADIO 3" | 134.89.45.24:30000
OA2 | 201104 | Terrace Point
| 201203 | Terrace Point
| 201304 | Terrace Point
| 201505 | Ano Nuevo | 2015-05-13T20:28:00Z | 2016-09-27T16:04:00Z |  | "OA2" | 134.89.45.24:20000 | ".72" written on radio can | JS inspected 12/2/2016, was labeled "OA RADIO 2" and IP 192.168.102.1, BLUE 468.9125. DataRadio by Calamp Viper400. Was found w/ no dessiccant, JS added.
| 201609 | Ano Nuevo | 2016-09-27T16:05:00Z | 2019-01-25T19:00:00Z | "OA2" | "OA RADIO 1" | 134.89.45.24:10000 | Drilled new bulkhead holes for VR2C and standard plug.
| 201812 | Ano Nuevo | 2019-01-25T21:00:00Z | 2020-07-16T16:00:00Z | "OA6" | "OA RADIO 3" | 134.89.45.24:30000
| 201912 | Ano Nuevo | 2020-08-13T18:10:00Z |  | "OA2" | "OA RADIO 1" | 134.89.45.24:10000

## CTD
Mooring Name | Deployment | Location | Inventory info
|  |  | tBeg | tEnd | Model | S/N | Comments
OA1 | 201203 | Hopkins
| 201306 | Hopkins |  |  | SBE-37SMP | 37SMP35635-3318
| 201401 | Hopkins | 2014-01-23T23:46:00Z | 2016-07-20T16:20:00Z | SBE-37SMP | 37SMP40327-4303
| 201607 | Hopkins | 2016-07-20T19:15:00Z | 2018-10-08T19:00:00Z | SBE-37SMP | 37SMP40327-3937 | Borrowed from J. Figurski. Tested in tank, data looked OK. Will do a post-recovery factory calibration. Has its own internal pump. Replaced batteries 7/19/2016.
| 201810 | Hopkins | 2018-11-16T10:30:00Z | 2020-10-20T19:00:00Z | SBE-37SMP | 37SMP35635-3318 | Low Salinity/Cond. Readings, Air vent clogged? Cleaned 10DEC2018. stopped working 20APR20. Calibrated at Seabird DEC20
| 202010 | Hopkins | 2020-10-20T23:25:00Z | 2022-03-04T18:00:00Z | SBE-37SMP | 37SMP417525-4590 | Borrowed from J. Figurski.
|  | Hopkins | 2022-03-04T18:00:00Z |  | SBE-37SMP | 37SMP35635-3318 | replaced due to memory loss
OA2 | 201104 | Terrace Point
| 201203 | Terrace Point
| 201304 | Terrace Point
| 201505 | Ano Nuevo | 2015-05-13T20:28:00Z | 2016-09-27T16:04:00Z | SBE37-SMP | 37SMP35635-3318 | Calibrated 2/2012. CW sent back for post-cal in Nov 2016.
| 201609 | Ano Nuevo | 2016-09-27T18:40:00Z | 2019-01-25T19:00:00Z | SBE-37SMP | 37SMP40327-4303 | Recovered from HOP OA1. Batteries were low. Replaced batteries 8/2/16. Tested in tank, data look good.
| 201812 | Ano Nuevo | 2019-01-25T21:00:00Z | 2020-07-16T16:00:00Z | SBE-37SMP | 37SMP40327-3937 | Damaged connector and conductivity cell--both fixed and calibrated on DEC20
| 201912 | Ano Nuevo | 2020-08-13T18:10:00Z |  | SBE-37SMP

## CO2
Mooring Name | Deployment | Location | Inventory info |  |  |  | Calibration info | [5 empty] | Standard | Stnd cal date | Comments
|  |  | tBeg | tEnd | Li-Cor S/N | System ID | Cal date | Cal file | Final pCO2 coefficients |  |  |  | ppm
[9 empty] | A | B | C | D
OA1 | 201104 | SandHill Bluff | 2011-04-13T16:40:00Z | 2011-11-16T18:16:00Z |  | OAM2 | 4/6/2011 | OAM_2.xlsm | -6.2875 | 0.1931 | 1.034 | -0.0010298
| 201203 | Hopkins
| 201306 | Hopkins |  |  |  | OA5 | 5/23/2013 | OA5_CO2Cal_2013.xlsx | -2.6114 | 0.2685 | 1.0068 | -0.0014696
| 201401 | Hopkins | 2014-01-23T23:46:00Z | 2016-07-20T16:20:00Z |  |  |  |  | -14.3968 | 0.4257 | 1.0147 | -0.00022448 | None | N/A | Did not use a standard this deployment.
| 201607 | Hopkins | 2016-07-20T19:15:00Z | 2017-03-22T17:59:00Z |  | OA5 | 6/29/2016 | 2016-06-29_OA5.cal | -9.5293 | 0.3627 | 0.9939 | 0.0009 | 482.54 | 7/18/2016 | Will set standard to sample every 7 days. Stopped reporting 1/24/2017 along with pCO2.
| 201607 | Hopkins | 2017-03-22T18:00:00Z | 2017-07-14T16:59:00Z |  | OA6 | 3/9/2017 | 2017-03-09_OA6.cal | -18.2665 | 1.0794 | 1.0038 | -0.0002 | None | 3/xx/2017 | For T-cal, used T > 8, CO2 < 1000 ppm. Can has stnd + VR2C. Std value 455.18, but leaked
| 201607 | Hopkins | 2017-07-14T17:00:00Z | 2017-10-11T14:50:00Z |  | OA6 | 3/9/2017 | 2017-03-09_OA6.cal | -18.2665 | 1.0794 | 1.0038 | -0.0002 | None | 6/26/2017 | Previous standard had a leak. Standard value 478.49, but leaked
| 201607 | Hopkins | 2017-10-11T17:00:00Z | 2018-03-28T12:00:00Z |  | OA5 | 9/11/2017 | 2017-09-11_OA5.cal | -27.1824 | 1.8109 | 0.9835 | 0.002 | 486.06 | 9/12/2017 | Previous stnd leaked again, suspect issue is at controller connector, swapping controllers.
| 201607 | Hopkins | 2018-03-28T12:00:00Z | 2018-10-08T19:00:00Z |  | OA5 | 9/11/2017 | 2017-09-11_OA5.cal | -27.1824 | 1.8109 | 0.9835 | 0.002 | 519.8 | 3/27/2018 | Leak in pCO2 plumbing, equilibrator, OA can and standard tank replaced on 3/28/2018
| 201810 | Hopkins | 2018-11-16T10:30:00Z | 2020-10-20T19:00:00Z |  | OA5 | 10/3/2018 | 2018-10-3_OA5 | -43.3345 | 2.5298 | 1.027 | 0.0001 | 519.75 | 11/xx/2018
| 202010 | Hopkins | 2020-10-20T23:25:00Z | 2021-06-10T16:00:00Z |  | OA6 | 2020-08-27 | 2020-08-27_OA6 | -21.7234 | 1.0111 | 1.029 | -0.0011026 | 449.7 | 10/2/2020 | for cal T<20, 10<CO2<900. rmse<1
| 202106 | Hopkins | 2021-06-10T20:00:00Z | 2022-01-27T22:00:00Z |  | OA2
| 202201 | Hopkins | 2022-01-27T22:00:00Z | 2022-03-04T18:00:00Z |  | OA6 | [6 empty] | 432.4 | 2022-01-26 | Batteries weren't chaging
| 202203 | Hopkins | 2022-03-04T18:00:00Z |  |  | OA2
OA2 | 201104 | Terrace Point |  |  |  | OA3 | xx/yy/2011 | soon_omega_08-11.xlsm | -22.4721 | 0.4665 | 1.093 | -0.0040903
| 201203 | Terrace Point
| 201304 | Terrace Point |  |  |  | OA2 | 4/2/2013 | 04-02-2013_Cal.xlsx | -11.5338 | 0.4636 | 1.0483 | -0.0027871
| 201505 | Ano Nuevo | 2015-05-12T10:28:00Z | 2016-09-27T16:04:00Z |  | OA6 | 4/23/2015 | 04-23-2015_OA6.cal | -25.2633 | 0.83015 | 1.0495 | -0.0019443 | None | N/A | Did not use a standard this deployment.
| 201609 | Ano Nuevo | 2016-09-27T18:40:00Z | 2019-01-25T19:00:00Z |  | OA2 | 8/17/2016 | 2016-08-17_OA2.cal | -1.8475 | 0.3633 | 1.0182 | -0.002 | 472.08 | 8/10/2016 | For T-cal, used 8.4 < T < 18.5 and 0 < xCO2 < 1000
| 201812 | Ano Nuevo | 2019-01-25T21:00:00Z | 2020-07-16T16:00:00Z |  | OA6 | 12/6/2018 | 2018-12-06_OA6.cal | -37.3486 | 1.699 | 1.0569 | -0.0024 | 485.66 | 1/2019
| 201912 | Ano Nuevo | 2020-08-13T18:10:00Z | 2021-04-01T18:10:00Z |  | OA2 | 11/5/2019 | 2019-11-05_OA2.cal | -0.0939 | 0.5944 | 1.0271 | -0.0016 | 483.8 | 6/2020 | 10<xCO2<900, all T (9 - 21.3C) , rmse < 1
| 201912 | Ano Nuevo | 2021-04-01T20:10:00Z |  |  | OA5 | 11/13/2020 | 2020-11-13_OA5.cal | -33.3843 | 1.6194 | 0.9868 | 0.0018 | 442.3 | 3/22/2021 | controller changed due to bad CF card? New coeffs taking CO2<900, rmse 1.6756. new co2 tank std#1

## pH
Mooring Name | Deployment | Location | Inventory info | [5 empty] | Calibration info | [5 empty] | Constants/Comments
|  |  | tBeg | tEnd | Logger Board | Cap Adapter | Electrode | Housing | Cal date | Batch # | pH of stnd | Vout | Final pH coefficients |  | R | F
[13 empty] | Ts | E0 | 8.3145 | 96487
OA1 | 201203 | Hopkins | 2012-03-13T00:00:00Z | 2012-06-07T00:00:00Z |  |  |  | M1 | 2012-03-23 | 99 | 7.907 | 0.04257 | 293.1 | -0.41727 | power failure 6/7/2012
| 201203 | Hopkins | 2012-08-01T00:00:00Z | 2012-10-17T00:00:00Z | N75162 | Y20000013918 | 1223Y200000014136 |  | 2012-07-31 |  |  |  | 294.44 | -0.39269 | internally-logged deployed as backup
| 201306 | Hopkins | 2013-06-25T15:14:00Z | 2014-01-17T17:29:00Z |  |  |  | M7 | 2013-06-13 | 126 | 8.0997 | 0.0616 | 287.21 | -0.39999
| 201401 | Hopkins | 2014-01-23T23:46:00Z | 2015-01-08T18:32:00Z | N/A |  |  | M5 | 2014-01-23 | 133 | 8.159 | 0.0849 | 280.66 | -0.36946
| 201401 | Hopkins | 2015-01-08T18:32:00Z | 2015-09-16T17:00:00Z | N/A |  |  | M6 | 2015-01-06 | 136 | 8.223 | 0.0905 | 283.44 | -0.37196
| 201401 | Hopkins | 2015-09-16T17:00:00Z | 2016-05-04T16:35:00Z | N/A |  |  | M7 | 2015-09-15 |  | 8.188 | 0.0795 | 279.13 | -0.37399
| 201401 | Hopkins | 2016-05-03T16:36:00Z | 2016-07-20T16:20:00Z | N/A |  |  | M3 | 2016-05-03 |  |  |  | 293.74 | -0.35184 | end cap is worn but still water tight
| 201607 | Hopkins | 2016-07-20T19:15:00Z | 2017-03-22T17:59:00Z | N/A | 1221Y200000013922 | 1336Y20000039711 | M8 | 2016-07-14 | 152 | 7.9274 | 0.08488 | 294.12 | -0.37776 | housing name may be wrong, or we named both M8? pH tailed off beginning 3/18/2017, had been out for tBeg + 8 months before tailing off.
| 201607 | Hopkins | 2017-03-22T18:00:00Z | 2017-07-28T18:29:00Z | N/A |  | 1336Y20000039714 | SC4 | 7/23/2015 | 144 | 8.0788 | 0.07793 | 284.99 | -0.37891 | Previously deployed at M1. Borrowed a new OCSD housing. No time to calibrate before deployment, but these cal's giving good values, agrees with pCO2w. Another cal was done earlier, don't know the date, V = 74.18, Ts = 21.05 C (needs calibration slope/int applied). Might try using these cal's when doing post-processing.
| 201607 | Hopkins | 2017-07-28T18:30:00Z | 2017-10-11T14:50:00Z | N/A |  | 1507Y200000062032 | SC1 | 6/26/2017 | 163 | 7.8877 | 0.05366 | 293.9659 | -0.40642 | Fresh electrode. Can SC1 is an OCSD (shorter, fatter style, fits OA1) can.
| 201607 | Hopkins | 2017-10-11T20:00:00Z | 2018-01-11T21:42:00Z | N/A | 1221Y200000013918 | 1507Y200000062032 | SC4 | 9/11/2017 | 163 | 7.89 | 0.06453 | 294.25 | -0.39613 | Previous pH sensor bad internal wiring, possibly battery disconnected on install?
| 201607 | Hopkins | 2018-01-11T21:43:00Z | 2018-03-28T16:00:00Z | N/A |  | 1507Y200000062026 | SC1 | 2018-01-03 | 169 | 7.8035 | 0.06576 | 293.6672 | -0.38894 | Previous pH bad readings, cable connector pins bent, ph sensor replaced on 1/11/18 at 21:43UTC
| 201607 | Hopkins | 2018-03-28T21:00:00Z | 2018-10-08T19:00:00Z | N/A | 1221Y200000013918 | 1507Y200000062032 | SC4 | 2018-03-18 | 171 | 7.9012 | 0.07108 | 295.1605 | -0.39166
| 201810 | Hopkins | 2018-11-16T10:30:00Z | 2020-10-20T19:00:00Z | N/A | N/A | 1507Y200000062026 | SC1 | 2018-11-14 | 177 | 7.833 | 0.06726 | 294.27 | -0.3901
| 202010 | Hopkins | 2020-10-20T23:25:00Z | 2021-06-10T16:00:00Z | N/A | N/A | 1808Y200000083061 | SC4 | 2020-09-04 | 187 | 7.9583 | 0.06388 | 293.19 | -0.39891 | new electrode
| 202106 | Hopkins | 2021-06-10T20:00:00Z | 2022-01-27T22:00:00Z
| 202201 | Hopkins | 2022-01-27T22:00:00Z |  | N/A | ?? | 2146Y200000094977 | M8 | 2022-01-26 | TRIS#6 |  | 0.08385 | 292.89 |  | new electrode
OA2 | 201104 | Terrace Point | 2011-04-11T00:00:00Z | 2011-08-30T00:00:00Z |  |  |  | M2 | 2011-04-06 | 99 | 7.972 | 0.0498 | 294.65 | -0.41628
| 201104 | Terrace Point | 2011-08-30T00:00:00Z | 2011-11-16T00:00:00Z |  |  |  | M4 | 2011-08-29 | 111 | 7.9 | 0.03215 | 293.09 | -0.42727
| 201203 | Terrace Point | 2012-03-13T00:00:00Z | 2012-08-14T00:00:00Z |  |  |  | M4 | 2012-03-21 | 111 | 7.895 | 0.03129 | 293.39 | -0.42831
| 201203 | Terrace Point | 2012-08-14T00:00:00Z | 2012-12-18T00:00:00Z | N44925 | Y189862900009 | 1222Y200000014041 |  | 2012-08-13 |  |  |  | 293.55 | -0.37568 | internally-logged
| 201304 | Terrace Point | 2013-04-23T00:00:00Z | 2013-11-07T00:00:00Z |  |  |  | M6 | 2013-04-22 | 126 | 8.159 | 0.08576 | 283.44 | -0.3731
| 201505 | Ano Nuevo | 2015-05-13T20:28:00Z | 2016-01-26T17:35:00Z | N/A | 1116Y191310800001 | 1511Y200000064477 | M3 | 2016-05-03 | 152 | 7.935 | 0.11061 | 293.74 | -0.35184 | All new parts. Original cal [5/6/2016; -0.332385, 278.67] did not agree as well with pCO2w as post-cal. Using post-cal.
| 201505 | Ano Nuevo | 2016-01-26T17:35:00Z | 2016-09-27T16:04:00Z | N/A | Y200000013918 | 1507Y200000062026 | M5 | 2016-01-14 | 144(2) | 7.932 | 0.07437 | 294.17 | -0.38861
| 201609 | Ano Nuevo | 2016-09-27T18:40:00Z | 2019-01-25T19:00:00Z | N/A | 1116Y191310800001 | 1511Y20000006447 | M8 | 8/23/2016 | 156 | 7.902 | 0.10607 | 294.33 | -0.35541 | Drilled new bulkhead plug hole in M8 can. Can has a ding in its side from a previous bottom-cluster deployment.
| 201812 | Ano Nuevo | 2019-01-25T21:00:00Z | 2020-07-16T16:00:00Z |  |  | 1507Y200000068233 | SC4 | 2018-12-20 | 177 | 7.815 | 0.05711 | 292.94 | -0.39933
| 201912 | Ano Nuevo | 2020-08-13T18:10:00Z |  |  |  | 1807Y200000082431 | No name | 25/11/2019 | 185 | 7.9408 | 0.05399 | 292.718 | -0.40723 | old style screw on cap

## O2
Mooring Name | Deployment | Location | Inventory info |  |  |  |  | Calibration info |  |  | Comments
|  |  | tBeg | tEnd | Model | S/N | Foil Batch | Cal date | Term1 | Term2
OA1 | 201203 | Hopkins
| 201306 | Hopkins
| 201401 | Hopkins | 2014-01-23T23:46:00Z | 2016-07-20T16:20:00Z | 4330 | 444
| 201607 | Hopkins | 2016-07-20T19:15:00Z | 2017-03-22T17:59:00Z | 4330 | 989 | 1206 |  |  |  | New Foil. No internal salinity compensation (Salinity = 0). Stopped reporting 1/24/2017 along with pCO2.
| 201607 | Hopkins | 2017-03-22T18:00:00Z | 2018-01-11T21:42:00Z | 4330A | 588 | ? |  |  |  | New Foil from new batch, JF calibrated 3/xx/2017. Optode failed during deployment and will be sent back to Aanderaa
| 201607 | Hopkins | 2018-01-11T21:43:00Z | 2018-10-08T19:00:00Z | 4330 | 442 | 1517 | ? |  |  | New Foil calibrated by Jules 1/2018? Installed on mooring 1/11/2018
| 201810 | Hopkins | 2018-11-16T10:30:00Z | 2020-10-20T19:00:00Z
| 202010 | Hopkins | 2020-10-20T23:25:00Z |  | 4330 | 442 | 1517 |  |  |  | New Foil calibrated by Jules in 2020
OA2 | 201104 | Terrace Point
| 201203 | Terrace Point
| 201304 | Terrace Point
| 201505 | Ano Nuevo | 5/13/2015 20:28 | 9/27/2016 16:04 | 4330A | 588
| 201609 | Ano Nuevo | 2016-09-27T18:40:00Z | 2019-01-25T19:00:00Z | 4330 | 444 | 1206 |  |  |  | New Foil, same batch as OA1.
| 201812 | Ano Nuevo | 2019-01-25T21:00:00Z | 2020-07-16T16:00:00Z
| 201912 | Ano Nuevo | 2020-08-13T18:10:00Z | [7 empty] | Ask Jules?

## Fluor
Mooring Name | Deployment | Location | Inventory info |  |  | Calibration info |  |  | Comments
|  |  | tBeg | tEnd | S/N | Cal date | CWO | SF
OA1 | 201203 | Hopkins
| 201306 | Hopkins
| 201401 | Hopkins | 2014-01-23T23:46:00Z | 2016-07-20T16:20:00Z | WS3S-555P | 11/7/2006 | 0.056 | 15.4 | Had poison inline. Calibration is way overdue.
| 201607 | Hopkins | 2016-07-20T19:15:00Z | 2018-10-08T19:00:00Z | WS3S-014P | 6/24/2016 | 0.072 | 8 | Previous factory cal had large negative value for SF, we sent it back for a quick re-cal.
| 201810 | Hopkins | 2018-11-16T10:30:00Z | 2020-10-20T19:00:00Z |  |  | 0.072 | 8 | Need Updating, Just used previous values
| 202010 | Hopkins | 2020-10-20T23:25:00Z |  | WS3S-014P | 6/24/2016 | 0.072 | 8 | Needs recalibrating--giving negative values
OA2 | 201104 | Terrace Point
| 201203 | Terrace Point
| 201304 | Terrace Point
| 201505 | Ano Nuevo | 2015-05-13T20:28:00Z | 2016-09-27T16:04:00Z | WS3S-441P | 10/27/2009 | 0.058 | 15 | Flow-through, old connector type.
| 201609 | Ano Nuevo | 2016-09-27T18:40:00Z | 2019-01-25T19:00:00Z | WS3S-555P | 11/7/2006 | 0.056 | 15.4 | Taken from OA1. Has poison inline. Calibration is way overdue.
| 201812 | Ano Nuevo | 2019-01-25T21:00:00Z | 2020-07-16T16:00:00Z | WS3S-014P | 6/24/2016 | 0.072 | 8 | updated after deployment
| 201912 | Ano Nuevo | 2020-08-13T18:10:00Z |  | WS3S-555P | 11/7/2006 | 0.056 | 15.4 | Old values, same serial number. Assume no factory cal unless known. Doesn't matter bc we want relative fluorescence anyways -- tested and responds to fluor strip

## MET
Mooring Name | Deployment | Location | Inventory info |  |  |  | Calibration info |  |  | Comments
|  |  | tBeg | tEnd | Model | S/N | Air Temp cal date | Air Temp slope | Air Temp int
OA1 | 201203 | Hopkins
| 201306 | Hopkins
| 201401 | Hopkins | 2014-01-23T23:46:00Z | 2016-07-20T16:20:00Z | PB-200 | 3201665 |  |  |  | S/N 3201665 was recovered 7/2016.
| 201607 | Hopkins | 2016-07-20T19:15:00Z | 2018-10-08T19:00:00Z | 200-WX | 3445950 | 6/22/2016 | 0.3401 | 9.4422 | Newer RS-232 style, more water-resistant. No humidity. Need to do a wind-dir calibration, instructions are in the manual. Air Temp cal regression derived from MLML Wx St'n while in parking lot.
| 201810 | Hopkins | 2018-11-16T10:30:00Z | 2020-10-20T19:00:00Z | 200-WX |  |  | 0.3401 | 9.4422 | Need updating, just used previous values
| 202010 | Hopkins | 2020-10-20T23:25:00Z |  | 200-WX | 60237809 | 10/16/2020 | 0.4123 | 8.212
OA2 | 201104 | Terrace Point
| 201203 | Terrace Point
| 201304 | Terrace Point
| 201505 | Ano Nuevo | 2015-05-13T20:28:00Z | 2016-01-26T17:35:00Z | [5 empty] | Flooded, fuse on OA controller blown.
| 201505 | Ano Nuevo | 2016-01-26T17:35:00Z | 2016-09-27T16:04:00Z | [5 empty] | Never reported any data
| 201609 | Ano Nuevo | 2016-09-27T18:40:00Z | 2017-09-27T17:50:00Z | 200-WX | 3445944 |  |  |  | Air T eventually went bad
| 201609 | Ano Nuevo | 2017-09-27T21:00:00Z | 2019-01-25T19:00:00Z | 200-WX | 3445949
| 201812 | Ano Nuevo | 2019-01-25T21:00:00Z | 2020-07-16T16:00:00Z |  |  |  | 0.3401 | 9.4422 | used same as OA1 to get T to plot, need update
| 201912 | Ano Nuevo | 2020-08-13T18:10:00Z |  | 200-WX | 4197256 | 8/2020 | 0.5177 | 6.5327 | cal over 3 months of parking lot data (longest to date)

## VR2C
Mooring Name | Deployment | Location | Inventory info |  |  |  |  | Comments
|  |  | tBeg | tEnd | Model | S/N | Frequency
OA1 | 201203 | Hopkins
| 201306 | Hopkins
| 201401 | Hopkins | 2014-01-23T23:46:00Z | 2016-07-20T16:20:00Z | VR2C | 450057 | 69 kHz | Had to cut the cable on recovery.
| 201607 | Hopkins | 2016-07-20T19:15:00Z | 2018-10-08T19:00:00Z | VR2C | 450093 | 69 kHz | Taken off wgOA.
| 201810 | Hopkins | 2018-11-16T10:30:00Z | 2020-10-20T19:00:00Z | VR2C | 450151 | 69 kHz
| 202010 | Hopkins | 2020-10-20T23:25:00Z |  | VR2C | 450093 | 69 kHz
OA2 | 201104 | Terrace Point
| 201203 | Terrace Point
| 201304 | Terrace Point
| 201505 | Ano Nuevo |  |  | VR2C | 450151 | 69 kHz
| 201609 | Ano Nuevo | 2016-09-27T18:40:00Z | 2019-01-25T19:00:00Z | VR2C | 450057 | 69 kHz
| 201812 | Ano Nuevo | 2019-01-25T21:00:00Z | 2020-07-16T16:00:00Z | VR2C | 450093 | 69 kHz
| 201912 | Ano Nuevo | 2020-08-13T18:10:00Z |  | VR2C | 450057 | 69 kHz