JS v.2017.06.26

1. Calibrations

Two types of calibrations are done for the pH sensors, one for the temperature measurement (self-contained units only) and one for pH (both ‘mooring style’ and self-contained units). Units that have an amplifier will also have a calibration for the amplifier. The A/D’s on these units do not need to be calibrated. The main purpose of the calibrations is to characterize the logger board; the electrodes and signal conditioning boards themselves theoretically don’t need to be characterized.

1.1. Temperature calibration

Temperature calibrations are typically done only for sensors that will NOT be deployed near a CTD. They are usually done when a new system, or new configuration of the same system (e.g., a new electrode only) will be used.

For the temperature calibrations, the electrode is kept at constant T and the resistance across the signal conditioning board’s voltage divider is measured with a multimeter. This is repeated for two additional temperatures, spanning the expected temperature range in the field, in order to check the linearity of the response. The resistance measurements are converted to temperature using ‘Steinhart’ constants:

SH_A = 1.06329736674527E-03

SH_B = 2.51377462346306E-04

SH_C = 2.55455247726963E-08

Temperature [K] = 1 ./ ( SH_A + (SH_B * log(Rtherm)) + (SH_C * (log(Rtherm).^3)) )

where Rtherm is the measured resistance in Ohms, and log is the natural logarithm1.

A linear, least-squares fit is then determined between the known and derived temperatures.

The result of the temperature calibration is three values: the battery voltage scale factor (VbatSF), and temperature slope and intercept.

1In practice, Rtherm is calculated from the logger board’s V2 (sealed positive battery voltage [-3.6 3.6] V), and V4 (Wheatstone bridge) outputs, and the signal conditioning board’s battery voltage scale factor (see pHcal.m for details).

1.2. pH calibration

The pH calibrations are done more often, typically every 6-12 months.

The pH calibrations use “certified reference materials (CRM) for the measurement of oceanic CO2 parameters” standards (Dickson, 2001). The CRM standard is poured into the electrode’s measurement chamber and the electrode’s voltage output is monitored with a multimeter. After a few minutes the liquid in the measurement chamber is poured out and replaced with fresh standard; this is repeated several times in order to get rid of any residue on the electrode. This is particularly important if the electrode had been in contact with a buffer solution.

The temperature of the standard is also continuously monitored, using a separate, calibrated thermistor. It may take up to 20 or 30 minutes for the temperature to stabilize. Calibrations are typically done in the lab, near room temperature.

Once these measurements are stable (electrode voltage within < 1 mV; temperature within 0.01 C), the values are recorded. The final result of the pH calibration is two values: Ts (the temperature of the CRM standard during calibration, in K); and E0, which is a function of the pH electrode voltage output (Vout), the known pH of the CRM standard (pHstd), and Ts:

E0 = Vout - (pHstd * R * Ts * log(10) / F), where:

R = 8.31451 (gas constant in J mol-1 K-1)

F = 96487 (converts between Coulombs and Faradays (1 F = 96487 C))

The CRM standards have known salinity, total alkalinity (TA in umol/kg), and total DIC (in umol/kg); pHstd is calculated from TA and DIC using the CO2SYS software package, inputting the temperature at calibration.

2. Processing of raw data

The logger board records four voltages, V1 through V4, and a time stamp. V1 is pH voltage, V2 is positive battery voltage, V3 is negative battery voltage, and V4 is temperature voltage. V2 and V4 are used to calculate temperature; V3 is not used. The conversions are performed in the Matlab functions calc_pH.m (when using T data from a nearby CTD), OR pHcal.m (when using the pH sensor’s internal temperature sensor). Please refer to these files for full details/comments and equations used.

Raw data are downloaded from the logger board using Madgetech software. Connect using the 3.5 mm audio jack to USB dongle and click the black and yellow down arrow icon. When the data are finished offloading, save the data to your hard drive using the ‘Save As…’ option. (The ‘Export’ option doesn’t include much header information, and formats the time stamp differently.)

Once the raw .csv file has been downloaded from the logger board, the data are processed by running pHcal.m. This function takes as inputs the raw .csv filename and directory, and the calibration coefficients obtained during the laboratory calibrations, and outputs converted pH [total scale] and temperature [C] values, which are written to a new .csv file with '*_conv.csv' in the filename.

3. Maintenance

The pH sensors have to be powered continuously. Ours are powered independently from the rest of the mooring. The electrode is powered by two 3.6 V AA lithium batteries, and the logger is powered by one 9 V lithium battery, both are inside the pH can. Fresh batteries should last 6-9 months. At a 15-min sampling interval the electrode batteries well outlast the logger batteries. An electrode battery failure will be evident in the data by a tapering-down of the main pH voltage (V1) over several hours to a couple days.

The quality of the pH data is entirely dependent on having a stable KCl gel as a reference. These gels can be depleted over time, though we have never seen one become depleted as long as they are kept in sea water (up to 2+ years and counting). Apparently there is a shorter lifetime for the KCl gel if the sensor is kept in fresh water as there would be a greater ion exchange between the gel and fresh water compared to salt water. Replacement gels can be ordered through Honeywell.

Between the measurement volume and the KCl gel reference is a white, plastic-looking barrier called a frit. The frit can dry out over time, compromising the division between the sample and reference and ruining the data. The frit needs to be kept in either a KCl solution (ideally), sea water, or at a minimum fresh water if the sensor is going to be out of the ocean for more than a few hours.

4. Bookkeeping

It is very important to keep all the sensor configuration and calibration information together and organized, including:

Logger board S/N (for self-contained units)

Electrode S/N

Cap adapter/amplifier S/N

Housing label

Dates of calibrations

CRM standard batch number

Temperature calibration values (VbatSF, slope, intercept)

pH calibration values (Vout, pHstd, Ts, E0)

Dates and locations/depths for each deployment

Notes for any re-configuration of the sensors, new batteries, failures, etc.