MBARI-CIMT Instrument Review

May 14, 2004

ECOBB BBSB-024R

ECOFL FLS-023

Instrument Reps: Mark Harlan, Reiko Michisaki

1. Known issues with the instrument configuraton and data streams: No issues known at this time.

2. Whether or not the instrument is ready to be deployed: Dark and light tests on sensors were not performed.

3. A brief description of the method used to QC the data stream: Raw data plotted and visually reviewed.  Calibrations are applied to the data and the calibrated data is plotted and visually reviewed.

4. Any pertinent configuration settings and their current settings: 
ECOBB calibrations are set as per calibration sheet from WetLabs.  
Scale factor for 660nm  = 3.72E-06
Angle  = 117
Wavelength = 660
Dark Counts = 63.870968
Cal SW = 0.00271
Salinity = 34.0
X = 1.1
ECOFL calibrations are set as per calibration sheet from WetLabs.  
Scale factor = 0.0079
CWO = 77

[image: image1.jpg]MBARI CIMT Mooring Backscatter (1301)

0.01

ECOBB Calibrated (10-4 m-1)

=t
|
|
I
1
|
|
I

e e a——

9
05/06

i
05/07

I Il I Il 1
05/08 05/09 05/10 05/11 05/12
Date

1500

1000

500

ECOBB Raw (counts)

9
05/06

I
05/07

I I i i
05/09 0510 05111 0512

Date

i
05/08




[image: image2.jpg]MBARI CIMT Mooring Fluorescence (1279)

e 2 2 2 a

o o o o o

& 0 & N &
T T T

=)

o

@
T

ECOFL Calibrated (mg m-3)

e

=
5N
=]
5

® ® ® ®
[ R )

@
RN

ECOFL Raw (counts)

81

i
I
1

]
T
[

i "\M‘“ /
i
iy

/
/





