## Info
Deployment | tBeg | tEnd
wgOA
20120316
20120521
20120829
20121015
20130123
20130315
20130506
20130625
20130909
20141217
20150825
20150923
wgTiny
20140318 | 2014-03-18T22:52:00Z | 2014-03-28T16:38:00Z
20140403 | 2014-04-11T21:30:00Z | 2014-04-21T23:46:00Z
20140521
20140619
20140918
20150212 | 2015-02-13T20:09:00Z | 2015-03-19T17:43:00Z
20150421 | 2015-04-22T22:44:00Z | 2015-04-28T16:44:00Z
20150501 | 2015-04-30T23:19:00Z | 2015-06-08T16:58:00Z
20150611
20150824
20150921 | 2015-09-21T18:46:00Z | 2015-10-14T19:26:00Z
20151019
20151209
20160128
20160225
20160315
20160627
20160701
20160830
20160920
20161005
20161031
20161212
20170109 | 2017-01-09T15:35:21Z | 2017-02-15T21:52:46Z
20170307 | 2017-03-07T18:00:00Z | 2017-04-05T00:20:00Z
20170412 | 2017-04-12T17:45:00Z | 2017-05-30T18:02:00Z
20170605 | 2017-06-05T18:45:00Z | 2017-06-08T17:51:00Z
20170630 | 2017-06-30T22:30:00Z | 2017-08-11T18:00:00Z
20170823 | 2017-08-23T19:30:00Z | 2017-09-12T18:40:00Z
20170921 | 2017-09-21T19:00:00Z | 2017-10-12T17:05:00Z
20171019 | 2017-10-19T21:30:00Z | 2017-12-11T20:00:00Z
20171212 | 2017-12-12T22:44:00Z | 2017-12-21T18:00:00Z
20180116 | 2018-01-16T23:19:00Z | 2018-02-01T17:38:00Z
20180208 | 2018-02-08T17:44:00Z | 2018-03-11T08:35:07Z
20180323 | 2018-03-23T17:20:00Z | 2018-05-01T18:00:00Z
20180516 | 2018-05-16T17:00:00Z | 2018-06-28T16:24:00Z
20180710 | 2018-07-10T22:00:00Z | 2018-07-17T21:20:00Z
20180730 | 2018-07-30T18:00:00Z | 2018-08-21T21:30:00Z
20180829 | 2018-08-29T21:00:00Z | 2018-09-21T20:25:00Z
20181001 | 2018-10-01T18:00:00Z | 2018-10-10T22:00:00Z
20181030 | 2018-10-30T18:00:00Z | 2018-12-03T22:00:00Z
20181210 | 2018-12-10T18:29:00Z | 2018-12-14T18:30:00Z
20181217 | 2018-12-17T19:30:00Z | 2018-12-20T19:00:00Z
20190114 | 2019-01-14T18:30:00Z | 2019-01-16T17:30:00Z
20190213 | 2019-02-13T20:00:00Z | 2019-02-19T20:00:00Z
20190304 | 2019-03-04T18:00:00Z | 2019-04-05T19:00:00Z
20190502 | 2019-05-02T22:50:00Z | 2019-05-09T19:30:00Z
20190513 | 2019-05-13T20:30:00Z | 2019-06-11T21:30:00Z
20190627 | 2019-06-27T00:15:00Z | 2019-07-10T19:00:00Z
20190719 | 2019-07-19T20:00:00Z | 2019-08-09T16:00:00Z
20190909 | 2019-09-09T22:30:00Z | 2019-09-17T16:30:00Z
20191001 | 2019-10-01T21:15:00Z | 2019-11-01T20:00:00Z
20191114 | 2019-11-14T20:00:00Z | 2019-11-15T21:00:00Z
20191213 | 2019-12-13T19:40:00Z | 2019-12-18T22:45:00Z
20200106 | 2020-01-06T18:30:00Z | 2020-01-10T17:30:00Z
20200203 | 2020-02-03T19:30:00Z | 2020-02-07T17:10:00Z
20200227 | 2020-02-27T20:15:00Z | 2020-03-05T21:45:00Z
20200708 | 2020-07-08T17:10:00Z | 2020-07-14T21:49:00Z
20200717 | 2020-07-17T23:00:00Z | 2020-09-01T22:30:00Z
20200903 | 2020-09-03T20:20:00Z | 2020-09-25T20:40:00Z
20201006 | 2020-10-06T00:00:00Z | 2020-11-17T00:15:00Z
20201202 | 2020-12-02T20:30:00Z | 2020-12-08T19:30:00Z
20210209 | 2021-02-09T20:10:00Z | 2021-02-18T19:00:00Z
20210318 | 2021-03-18T18:40:00Z | 2021-04-02T23:00:00Z
20210408 | 2021-04-08T19:40:00Z | 2021-05-14T17:00:00Z
20210528 | 2021-05-28T18:56:00Z | 2021-06-28T19:46:00Z
20210701 | 2021-07-01T19:00:00Z | 2021-07-08T17:00:00Z
20210804 | 2021-08-04T17:43:00Z | 2021-08-10T18:00:00Z
20210818 | 2021-08-18T17:12:00Z | 2021-08-25T18:16:00Z
20210902 | 2021-09-02T18:06:00Z | 2021-09-10T16:51:00Z
20210928 | 2021-09-28T17:09:00Z | 2021-10-20T17:41:00Z
20211028 | 2021-10-28T19:26:00Z | 2021-11-09T21:30:00Z
20211129 | 2021-11-29T20:10:00Z | 2021-12-06T23:04:00Z
20211215 | 2021-12-15T19:10:00Z | 2021-12-17T19:55:00Z
20220107 | 2022-01-07T18:04:00Z | 2022-01-13T18:28:00Z
20220124 | 2022-01-24T21:10:00Z | 2022-02-23T21:19:00Z
20220310 | 2022-03-10T18:33:00Z | 2022-04-14T17:39:00Z
20220420 | 2022-04-20T21:47:00Z | 2022-05-12T16:53:00Z
20220516 | 2022-05-16T19:08:00Z | 2022-05-18T19:58:00Z
20220524 | 2022-05-24T01:15:00Z | 2022-06-01T19:43:00Z
20220602 | 2022-06-02T22:59:00Z | 2022-06-07T20:12:00Z
20220616 | 2022-06-16T18:33:00Z | 2022-07-07T16:58:00Z
20220719 | 2022-07-19T18:27:00Z | 2022-08-18T18:37:00Z
20220909 | 2022-09-09T17:06:00Z | 2022-09-15T18:21:00Z
20220919 | 2022-09-19T18:43:00Z | 2022-10-12T17:48:00Z
20221028 | 2022-10-28T18:20:00Z | 2022-10-28T23:25:00Z
20221103 | 2022-11-03T18:50:00Z | 2022-11-22T18:15:00Z
20221209 | 2022-12-09T17:48:00Z | 2022-12-09T23:31:00Z
20221214 | 2022-12-14T22:28:00Z | 2022-12-19T20:43:00Z
20230118 | 2023-01-18T21:15:00Z | 2023-01-24T18:50:00Z
20230215 | 2023-02-15T22:22:00Z | 2023-02-21T17:27:00Z
20230302 | 2023-03-02T18:05:00Z | 2023-03-16T16:18:00Z
20230406 | 2023-04-06T17:34:00Z | 2023-05-04T18:44:00Z
20230512 | 2023-05-12T17:35:00Z | 2023-05-30T16:39:00Z
20230605 | 2023-06-05T18:26:00Z | 2023-06-15T17:21:00Z
20230710 | 2023-07-10T19:46:00Z | 2023-07-19T20:41:00Z
20230731 | 2023-07-31T22:50:00Z | 2023-08-18T20:16:00Z
20230908 | 2023-09-08T18:15:00Z | 2023-09-13T17:00:00Z
20231013 | 2023-10-13T17:45:00Z | 2023-10-17T17:50:00Z
20231114 | 2023-11-14T18:50:00Z | 2023-11-17T20:01:00Z
20231208 | 2023-12-08T19:43:00Z | 2023-12-13T20:53:00Z
20231220 | 2023-12-20T18:41:00Z | 2023-12-22T20:05:00Z
20240123 | 2024-01-23T23:49:00Z | 2024-01-30T17:34:00Z
20240215 | 2024-02-15T22:26:00Z | 2024-02-22T18:06:00Z
20240311 | 2024-03-11T18:46:00Z | 2024-03-15T17:54:00Z
20240418 | 2024-04-18T19:42:00Z | 2024-05-01T18:22:00Z
20240522 | 2024-05-22T18:15:00Z | 2024-06-04T18:39:00Z
20240618 | 2024-06-18T21:13:00Z | 2024-06-25T17:00:00Z
20240719 | 2024-07-19T18:26:00Z | 2024-07-23T19:40:00Z
20240813 | 2024-08-13T20:54:00Z | 2024-08-16T17:35:00Z
20240911 | 2024-09-11T19:19:00Z | 2024-09-30T17:41:00Z
20241024 | 2024-10-24T16:46:00Z | 2024-10-31T18:23:00Z
20241108 | 2024-11-08T18:24:00Z | 2024-11-14T19:25:00Z
20241216 | 2024-12-16T19:04:00Z | 2024-12-20T19:00:00Z
20250107 | 2025-01-07T20:15:00Z | 2025-01-14T20:45:00Z
20250131 | 2025-01-31T17:52:00Z | 2025-02-03T19:21:00Z
20250403 | 2025-04-03T21:45:00Z | 2025-04-09T17:52:00Z
20250514 | 2025-05-14T18:36:00Z | 2025-05-23T19:05:00Z
20250606 | 2025-06-06T19:39:00Z | 2025-06-17T18:50:00Z
20250717 | 2025-07-17T16:14:00Z | 2025-07-28T18:13:00Z
20250813 | 2025-08-13T17:51:00Z | 2025-08-20T18:26:00Z
20250904 | 2025-09-04T19:47:00Z | 2025-10-08T19:38:00Z
20251020 | 2025-10-20T18:16:00Z | 2025-10-29T19:08:00Z
wgSparky
20160523 | 2016-05-23T17:05:00Z | 2016-06-30T22:58:00Z
20160630
20180531 | 2018-05-31T17:30:00Z | 2018-06-15T19:00:00Z
20180821 | 2018-08-21T21:30:00Z | 2018-08-27T16:10:00Z
20180905 | 2018-09-05T22:00:00Z | 2018-09-18T16:50:00Z
20181106 | 2018-11-06T19:00:00Z | 2018-11-08T19:00:00Z
20181113 | 2018-11-13T23:59:00Z | 2018-11-20T16:30:00Z
wgHansen
20190415 | 2019-04-15T21:00:00Z | 2019-04-29T20:30:00Z
20190522 | 2019-05-22T17:30:00Z | 2019-06-12T17:30:00Z
20190625 | 2019-06-25T21:00:00Z | 2019-07-17T18:00:00Z
20190812 | 2019-08-12T18:00:00Z | 2019-09-03T17:00:00Z
20190927 | 2019-09-27T18:00:00Z | 2019-10-16T18:00:00Z
20191209 | 2019-12-09T23:45:00Z | 2019-12-17T19:00:00Z
20191218 | 2019-12-18T21:00:00Z | 2019-12-19T23:00:00Z
20200619 | 2020-06-19T18:20:00Z | 2020-07-08T17:10:00Z
20200716 | 2020-07-16T23:10:00Z | 2020-08-05T18:00:00Z
20200909 | 2020-09-09T20:10:00Z | 2020-09-28T17:30:00Z
20201005 | 2020-10-05T18:40:00Z | 2020-10-26T18:10:00Z
20201113 | 2020-11-13T20:50:00Z | 2020-11-24T23:35:00Z
20201204 | 2020-12-04T19:00:00Z | 2020-12-10T19:30:00Z
20210218 | 2021-02-18T19:20:00Z | 2021-03-17T18:40:00Z
20210319 | 2021-03-19T00:10:00Z | 2021-04-06T18:10:00Z
20210409 | 2021-04-09T17:00:00Z | 2021-04-27T17:20:00Z
20210518 | 2021-05-18T17:55:00Z | 2021-06-07T16:58:00Z
20210625 | 2021-06-25T18:55:00Z | 2021-07-01T18:48:00Z
20210810 | 2021-08-10T18:00:00Z | 2021-08-30T19:47:00Z
20211026 | 2021-10-26T18:18:00Z | 2021-11-15T22:35:00Z
20220118 | 2022-01-18T21:44:00Z | 2022-01-20T17:53:00Z
20220202 | 2022-02-02T00:37:00Z | 2022-02-04T17:58:00Z
20220209 | 2022-02-09T18:37:00Z | 2022-02-10T21:47:00Z
20220216 | 2022-02-16T18:17:00Z | 2022-03-07T16:17:00Z
20220323 | 2022-03-23T17:49:00Z | 2022-04-11T17:30:00Z
20220420 | 2022-04-20T20:38:00Z | 2022-05-10T18:18:00Z
20220525 | 2022-05-25T18:20:00Z | 2022-06-15T17:45:00Z
20220705 | 2022-07-05T14:48:00Z | 2022-07-27T19:38:00Z
20220815 | 2022-08-15T22:55:00Z | 2022-09-07T17:55:00Z
20220908 | 2022-09-08T17:11:00Z | 2022-09-28T16:46:00Z
20220929 | 2022-09-29T16:39:00Z | 2022-10-18T18:14:00Z
20221020 | 2022-10-20T17:00:00Z | 2022-11-07T18:25:00Z
20221109 | 2022-11-09T21:07:00Z | 2022-11-29T18:42:00Z
20221130 | 2022-11-30T23:21:00Z | 2022-12-20T19:21:00Z
20230112 | 2023-01-12T18:46:00Z | 2023-02-07T17:56:00Z
20230209 | 2023-02-09T18:58:00Z | 2023-03-02T18:11:00Z
20230404 | 2023-04-04T16:07:00Z | 2023-04-24T18:40:00Z
20230515 | 2023-05-15T18:43:00Z | 2023-05-22T19:08:00Z
20230607 | 2023-06-07T18:46:00Z | 2023-06-26T18:53:00Z
20230707 | 2023-07-07T17:24:00Z | 2023-07-26T17:43:00Z
20230809 | 2023-08-09T18:07:00Z | 2023-08-30T16:53:00Z
20230901 | 2023-09-01T17:24:00Z | 2023-09-21T18:54:00Z
20230927 | 2023-09-27T19:18:00Z | 2023-10-16T19:43:00Z
20231018 | 2023-10-18T17:13:00Z | 2023-10-25T16:55:00Z
20231031 | 2023-10-31T16:53:00Z | 2023-11-01T17:45:00Z
20231108 | 2023-11-08T18:41:00Z | 2023-11-29T00:26:00Z
20231130 | 2023-11-30T16:54:00Z | 2023-12-19T17:34:00Z
20240118 | 2024-01-18T16:20:00Z | 2024-02-21T18:24:00Z
20240228 | 2024-02-28T20:19:00Z | 2024-03-20T17:58:00Z
20240415 | 2024-04-15T18:57:00Z | 2024-05-06T18:22:00Z
20240516 | 2024-05-16T20:28:00Z | 2024-06-06T16:45:00Z
20240708 | 2024-07-08T18:11:00Z | 2024-07-29T21:50:00Z
20240807 | 2024-08-07T18:22:00Z | 2024-08-28T17:12:00Z
20240911 | 2024-09-11T17:06:00Z | 2024-10-02T17:21:00Z
20241007 | 2024-10-07T20:55:00Z | 2024-10-28T18:08:00Z
20241119 | 2024-11-19T18:35:00Z | 2024-12-10T19:02:00Z
20250107 | 2025-01-07T18:15:00Z | 2025-01-27T17:59:00Z
20250311 | 2025-03-11T17:16:00Z | 2025-04-07T22:14:00Z
20250415 | 2025-04-15T22:57:00Z | 2025-05-05T18:38:00Z
20250527 | 2025-05-27T19:58:00Z | 2025-06-16T18:36:00Z
20250709 | 2025-07-09T21:35:00Z | 2025-07-29T20:01:00Z
20250820 | 2025-08-20T17:32:00Z | 2025-09-09T19:43:00Z
20251001 | 2025-10-01T19:07:00Z | 2025-10-17T19:01:00Z
wgTest
20230509 | 2023-05-09T18:59:00Z | 2023-05-12T17:37:00Z
20231129 | 2023-11-29T21:51:00Z | 2023-11-30T17:48:00Z
20240304 | 2024-03-04T20:13:00Z | 2024-03-06T22:00:00Z
20240321 | 2024-03-21T17:14:00Z | 2024-03-22T16:34:00Z
20250604 | 2025-06-04T17:35:00Z | 2025-06-11T17:03:00Z

## CCU
Deployment | CCU | VID | Notes
wgOA
20120316
20120521
20120829
20121015
20130123
20130315
20130506
20130625
20130909
20141217
20150825
20150923
wgTiny
20140318 | Tiny | 794595959
20140403 | Tiny | 794595959
20140521 | Tiny | 794595959
20140619 | Tiny | 794595959
20140918 | Tiny | 794595959
20150212 | Tiny | 794595959
20150421 | Tiny | 794595959
20150501 | Tiny | 794595959
20150611 | Tiny | 794595959
20150824 | Tiny | 794595959
20150921 | Tiny | 794595959
20151019 | Tiny | 794595959
20151209 | TinyToo | 1531399927
20160128 | TinyToo | 1531399927
20160225 | TinyToo | 1531399927
20160315 | Tiny | 794595959
20160627 | Tiny | 794595959
20160701 | Tiny | 794595959
20160830 | TinyToo | 1531399927
20160920 | TinyToo | 1531399927
20161005 | TinyToo | 1531399927
20161031 | TinyToo | 1531399927
20161212 | TinyToo | 1531399927
20170109 | TinyToo | 1531399927
20170307 | TinyToo | 1531399927
20170412 | TinyToo | 1531399927
20170605 | TinyToo | 1531399927
20170630 | TinyToo | 1531399927
20170823 | TinyToo | 1531399927
20170921 | TinyToo | 1531399927
20171019 | TinyToo | 1531399927
20171212 | TinyToo | 1531399927
20180116 | Tiny | 794595959
20180208 | Tiny | 794595959
20180323 | Tiny | 794595959
20180516 | Tiny | 794595959
20180710 | Tiny | 794595959
20180730 | Tiny | 794595959
20180829 | Tiny | 794595959
20181001 | Tiny | 794595959
20181030 | Tiny | 794595959
20181210 | Tiny | 794595959
20181217 | Tiny | 794595959
20190114 | Tiny | 794595959
20190213 | Tiny | 794595959
20190304 | Tiny | 794595959
20190502 | Tiny | 794595959
20190513 | Tiny | 794595959
20190627 | Tiny | 794595959
20190719 | MBARI Spare | 1531399927 | CCU Tinytoo was renamed MBARI Spare
20190909 | MBARI Spare | 1531399927 | CCU Tinytoo was renamed MBARI Spare
20191001 | MBARI Spare | 1531399927 | CCU Tinytoo was renamed MBARI Spare
20191114 | Hansen | 1553177065 | Hansens CCU was used on vehicle Tiny for this deployment
20191213 | Hansen | 1553177065 | Hansens CCU was used on vehicle Tiny for this deployment
20200106 | Hansen | 1553177065 | Hansens CCU was used on vehicle Tiny for this deployment
20200203 | Tiny | 794595959
20200227 | Tiny | 794595959
20200708 | MBARI Spare | 1531399927
20200717 | MBARI Spare | 1531399927 | CCU shipped back to LR
20200903 | Tiny | 794595959
20201006 | Tiny | 794595959
20201202 | Tiny | 794595959
20210209 | Tiny | 794595959
20210318 | Tiny | 794595959
20210408 | Tiny | 794595959
20210528 | Tiny | 794595959 | This excel spreadsheet was left open so I couldn't set up processing until 02JUN. Missing data was dowloaded on 29JUN
20210701 | Tiny | 794595959
20210804 | Tiny | 794595959
20210818 | Tiny | 794595959
20210902 | Tiny | 794595959 | Wave Glider Tiny was attacked by a shark during this mission and suffered minor damage.
20210928 | Tiny | 794595959 | Wave Glider Tiny's hull was repaired before this mission!
20211028 | Tiny | 794595959
20211129 | Tiny | 794595959
20211215 | Tiny | 794595959 | Plots were started after deployment, data from this mission needs to be downloaded but the script still appears to be broken…
20220107 | Tiny | 794595959
20220124 | Tiny | 794595959 | Data gap on 26JAN2022, download data after deployment
20220310 | Tiny | 794595959
20220420 | Tiny | 794595959 | Plots started after deployment download complete dataset after deloyment
20220516 | Tiny | 794595959
20220524 | Tiny | 794595959 | ECOHAB mission, deployed and recovered by Brian K. off the Goleta Pier!
20220602 | Tiny | 794595959
20220616 | Tiny | 794595959
20220719 | Tiny | 794595959 | Tinys orig. hull damaged on this deployment.
20220909 | Tiny | 794595959 | JRF spare hull installed.
20220919 | Tiny | 794595959
20221028 | Tiny | 794595959
20221103 | Hansen Replacement | 466563798 | Power failure on WG Tiny's CCU before last deployment and a new CCU from LR "Hansen's replacement" was used for this deployment. Plots were started on 14NOV2022 so download data after deployment.
20221209 | Hansen Replacement | 466563798
20221214 | Hansen Replacement | 466563798
20230118 | Hansen Replacement | 466563798 | Wave Glider Tiny was serverly fouled in kelp during this deployment between m1 and m2. The vehicle was disabled by the kelp, drifted south and was recovered ~12nm offshore of Point Sur. Umbilical was very twistd when we recovered. There were float to sub comms errors and also a ground fault. Umbilical was replaced with Tiny's Original v200 umblical which was repaired under RMA ### for bushing spacing. replacement used umbical had about 200 to 300 days at sea on it already
20230215 | Hansen Replacement | 466563798 | Umbilical replaced with a spare used one, MF DAT cable also replaced.
20230302 | Hansen Replacement | 466563798
20230406 | Hansen Replacement | 466563798
20230512 | Hansen Replacement | 466563798
20230605 | Hansen Replacement | 466563798
20230710 | Hansen Replacement | 466563798
20230731 | Hansen Replacement | 466563798
20230908 | Hansen Replacement | 466563798
20231013 | Hansen Replacement | 466563798
20231114 | Hansen Replacement | 466563798
20231208 | Hansen Replacement | 466563798
20231220 | Hansen Replacement | 466563798 | WetWiFI engineering tests. Science sensors not run for this testing. pCO2 and pH removed for this deployment.
20240123 | Hansen Replacement | 466563798
20240215 | Hansen Replacement | 466563798
20240311 | Hansen Replacement | 466563798
20240418 | SV3-110 Wave Glider Tiny | 466563798 | Changed CCU SV3-110s name to Wave Glider Tiny
20240522 | SV3-110 Wave Glider Tiny | 466563798
20240618 | SV3-110 Wave Glider Tiny | 466563798
20240719 | SV3-110 Wave Glider Tiny | 466563798
20240813 | SV3-110 Wave Glider Tiny | 466563798
20240911 | SV3-110 Wave Glider Tiny | 466563798
20241024 | SV3-110 Wave Glider Tiny | 466563798
20241108 | SV3-110 Wave Glider Tiny | 466563798
20241216 | SV3-110 Wave Glider Tiny | 466563798
20250107 | SV3-110 Wave Glider Tiny | 466563798
20250131 | SV3-110 Wave Glider Tiny | 466563798
20250403 | SV3-110 Wave Glider Tiny | 466563798
20250514 | SV3-110 Wave Glider Tiny | 466563798
20250606 | SV3-110 Wave Glider Tiny | 466563798
20250717 | SV3-110 Wave Glider Tiny | 466563798
20250813 | SV3-110 Wave Glider Tiny | 466563798
20250904 | SV3-110 Wave Glider Tiny | 466563798
20251020 | SV3-110 Wave Glider Tiny | 466563798 | Plots/processing started late will need a download afterwards
wgSparky
20160523 | Sparky | 199056737
20160630 | Sparky | 199056737
20180531 | Sparky | 199056737
20180821 | Sparky | 199056737
20180905 | Sparky | 199056737
20181106 | Sparky | 199056737
20181113 | Sparky | 199056737 | Vehicle Sparky was damged and lost during this deployment
wgHansen
20190415 | Hansen | 1553177065 | Shakedown deployment
20190522 | Hansen | 1553177065 | Acoustic Deployment for CANON19S
20190625 | Hansen | 1553177065
20190812 | Hansen | 1553177065
20190927 | Hansen | 1553177065
20191209 | MBARI Spare | 1531399927 | LRAUV ESP tracking
20191218 | MBARI Spare | 1531399927 | AUV tow testing
20200619 | Hansen | 1553177065 | KBB acoustics
20200716 | Hansen | 1553177065 | CANON & KBB Acoutics
20200909 | Hansen | 1553177065 | Acoustics/Shipless CANON
20201005 | Hansen | 1553177065 | Acoustics/Shipless CANON
20201113 | Hansen | 1553177065 | Acoustics
20201004 | MBARI Spare | 1531399927 | Acoustics make-up, different CCU
20210218 | MBARI Spare | 1531399927 | Acoustics
20210319 | MBARI Spare | 1531399927 | Astrid's acoustics
20210409 | MBARI Spare | 1531399927 | CANON Spring
20210518 | MBARI Spare | 1531399927 | Acoustics
20210625 | MBARI Spare | 1531399927 | Hotspot test, plots not started until 29jun, complete mission data downloaded on 17AUG please check data files
20210810 | SV3-029 | 1167408762 | Bioacoustics mission, had to run SV3-029 due to issues with two of MBARI's CCUs, float reboots on 22AUG, the vehicle was recovered on 25AUG, inspected on the deck of the Paragon and redeployed. CCU was not on MBARI Data Portal for the deployment, it was added by LR and mission data was downloaded on 24MAR2022, please check data file!
20211026 | MBARI Spare | 1531399927 | Processing/plotting started on 28OCT download data for entire deployment once misison is complete!
20220118 | MBARI Spare | 1531399927 | Winch test deployment
20220202 | MBARI Spare | 1531399927 | WetWiFi test deployment
20220209 | MBARI Spare | 1531399927 | KBB Transducer test
20220216 | MBARI Spare | 1531399927 | KBB acoutics mission, new 2022 route in upper MB canyon running new transducer setup
20220323 | MBARI Spare | 1531399927 | KBB acoutics mission, Astrid short transect line between m1 and point Pinos
20220420 | MBARI Spare | 1531399927 | KBB acoutics mission, new 2022 route in upper MB canyon running new transducer setup, plots started on 27APR make sure to download complete dataset afterwards.
20220525 | MBARI Spare | 1531399927 | AyeRIS/Ivan testing
20220705 | MBARI Spare | 1531399927
20220815 | MBARI Spare | 1531399927
20220908 | MBARI Spare | 1531399927
20220929 | MBARI Spare | 1531399927
20221020 | MBARI Spare | 1531399927
20221109 | MBARI Spare | 1531399927
20221130 | MBARI Spare | 1531399927
20230112 | MBARI Spare | 1531399927
20230209 | MBARI Spare | 1531399927
20230404 | MBARI Spare | 1531399927
20230515 | SV3-029 | 1167408762 | MBARI spare CCU SV3-121 leaked before deployment and SV3-029 was installed.
20230607 | SV3-029 | 1167408762
20230707 | SV3-029 | 1167408762
20230809 | SV3-029 | 1167408762
20230901 | SV3-029 | 1167408762
20230927 | SV3-029 | 1167408762
20231018 | SV3-029 | 1167408762 | Canbus and Nav issues during this deployment and the vehicle was recovered early.
20231031 | SV3-006 | 1338599832 | CCU replaced due to bulkhead connector failure. Leaking on Canbus boards? Weird nav issues on this deployment also, but vehicle fixed itself when back onshore.
20231108 | SV3-006 | 1338599832
20231130 | SV3-006 | 1338599832
20240118 | SV3-006 | 1338599832
20240228 | SV3-006 | 1338599832 | S4 bulkhead failure, Jordan update the software on SV3-034 and installed the payload.
20240415 | SV3-034 | 16977098 | Double check VID!
20240516 | SV3-034 | 16977098
20240708 | SV3-034 | 16977098
20240807 | SV3-034 | 16977098
20240911 | SV3-034 | 16977098
20241007 | SV3-034 | 16977098
20241119 | SV3-034 | 16977098
20250107 | SV3-034 | 16977098 | Plots/processing started late will need a download afterwards
20250311 | SV3-034 | 16977098 | Plots/processing started late will need a download afterwards
20250415 | SV3-034 | 16977098 | Plots/processing started late will need a download afterwards
20250527 | SV3-034 | 16977098 | Plots/processing started late will need a download afterwards
20250709 | SV3-034 | 16977098 | Plots/processing started late will need a download afterwards
20250820 | SV3-038 | 209609080 | CCU SV3-034 was not holding full pressure. CCU SV3-038 was brought up, tested and installed.
20251001 | SV3-038 | 209609080
wgTest
20230509 | SV3-029 | 1167408762 | Test Wave Glider built and used for Front Tracking testing
20231129 | SV3-172 | 1993437788 | compass issues and umblical damged from twisting
20240304 | SV3-172 | 1993437788 | thrudder replaced and vehicle appears to be operational!
20240321 | SV3-172 | 1993437788
20250604 | SV3-172 | 1993437788

## CO2
Deployment | Inventory info |  | Calibration info | [12 empty] | Standard | Stnd cal date | Comments | 5/5/2014 cal data are in: Diatom/WG/DD/2014_May_SV3/05-05-2014_CO2_TEMP.CAL
| Li-Cor S/N | System ID | Cal date | Final pCO2 coefficients | [11 empty] | ppm |  |  | 10/x/2012 cal data are in: Diatom/WG/WG_cold_cal.xlsx
wgOA |  |  |  | A | B | C | D | E | F | G | H | I | J | K | L
20120316 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284 | N/A | N/A | pCO2 data looks like all garbage.
20120521 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284 | N/A | N/A
20120829 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284 | 422.56 | 7/10/2012
20121015 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284 | N/A | N/A
20130123 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284; Final pCO2 coefficients E: 0.997 | 428.87 | 1/x/2013
20130315 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284 | 428.87 | 3/7/2013
20130506 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284 | 401.16 | 4/29/2013
20130625 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284 | 401.2 | N/A
20130909 | N/A | WG1 | 10/x/2012 | Final pCO2 coefficients A: -16.5956; Final pCO2 coefficients B: 0.3055; Final pCO2 coefficients C: 1.0901; Final pCO2 coefficients D: -0.00284; Final pCO2 coefficients E: 0.98946 | 401.2 | N/A
20141217 | N/A | WG2 | 9/9/2015 | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A
20150825 | N/A | WG2 | 9/9/2015 | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A
20150923 | N/A | WG2 | 9/9/2015 | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | The CO2 payload flooded, SV2 was then re-deployed 10/1/2015 without the CO2 payload under this same deployment name.
wgTiny
20140318 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | N/A | N/A | NO PCO2 DATA ON FETCH. MAY BE RAW PCO2 DATA ON CF.
20140403 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 398.22 | 4/30/2014 | NO PCO2 DATA ON FETCH. RAW PCO2 DATA ON CF ONLY.
20140521 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 430 | 5/12/2014
20140619 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 430 | 5/12/2014
20140918 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 429.93 | 11/6/2014 | An additional post-cal was performed on 11/12/14 (427.77)
20150212 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 427.41 | 4/9/2015
20150421 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 408.29 | 4/16/2015 | Stnd Cal measured and entered by CDW on 15APR2015 DP was -57.76 deg C
20150501 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 408.29 | 4/16/2015
20150611 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 408.29 | 4/16/2015 | Payload 3 (OA) went out after first few da
20150824 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -9.7075; Final pCO2 coefficients B: 0.4544; Final pCO2 coefficients C: 1.0281; Final pCO2 coefficients D: -0.0014; Final pCO2 coefficients E: 0.99515 | 406.17 | 08/20/2015 | New Standard. Custom cal process based on comparison w/ underway Carson, for this depl only.
20150921 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 406.17 | 08/20/2015
20151019 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 406.17 | 08/20/2015 | No OA data due to CTD payload failure
20151209 | N/A | WG1 | 5/5/2014 | -5.4872e-05 | 0.030826 | -1.7203 | -9.9368e-08 | 6.2298e-05 | -0.058365 | 5.2634e-06 | -0.0020836 | 1.703 | -6.1803e-05 | 1.0173 | -11.505 | 406.46 | N/A | Standard read ~411 ppm throughout deployment
20160128 | N/A | WG1 | 5/5/2014 | -5.4872e-05 | 0.030826 | -1.7203 | -9.9368e-08 | 6.2298e-05 | -0.058365 | 5.2634e-06 | -0.0020836 | 1.703 | -6.1803e-05 | 1.0173 | -11.505 | 406.46 | N/A | Post-cal on SV3 pCO2 standard tank. Standard read ~411 ppm throughout deployment. No CTD data, so pCO2 temperatures used in derivation of pCO2w and pH.
20160225 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | N/A | N/A | No science data collected this deployment
20160315 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | N/A | N/A | Standard read ~471 ppm for most of deployment, then ran out (~409 ppm, atmospheric).
20160627 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | N/A | N/A | Standard tank empty throughout deployment
20160701 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | N/A | N/A | Standard tank empty throughout deployment
20160830 | N/A | WG1 | 5/5/2014 | -5.4872e-05 | 0.030826 | -1.7203 | -9.9368e-08 | 6.2298e-05 | -0.058365 | 5.2634e-06 | -0.0020836 | 1.703 | -6.1803e-05 | 1.0173 | -11.505 | 467.58 | N/A | Standard re-filled prior to deployment
20160920 | N/A | WG1 | 5/5/2014 | -5.4872e-05 | 0.030826 | -1.7203 | -9.9368e-08 | 6.2298e-05 | -0.058365 | 5.2634e-06 | -0.0020836 | 1.703 | -6.1803e-05 | 1.0173 | -11.505 | 467.58 | N/A
20161005 | N/A | WG1 | 5/5/2014 | -5.4872e-05 | 0.030826 | -1.7203 | -9.9368e-08 | 6.2298e-05 | -0.058365 | 5.2634e-06 | -0.0020836 | 1.703 | -6.1803e-05 | 1.0173 | -11.505 | 467.58 | N/A
20161031 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 467.58 | N/A | Not implementing 'new' CO2 cal-correction per conversation with Gernot.
20161212 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 467.58 | N/A
20170109 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 467.58 | N/A
20170307 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 467.58 | N/A
20170412 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 546.18 | N/A | Re-filled with higher conc.
20170605 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 546.18 | N/A
20170630 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 546.18 | N/A
20170823 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 546.18 | N/A
20170921 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 460.4 | N/A | New stnd, wasn't cal'd prior to deployment. There is a note on standard tank from 9/18/17 claiming and 380 ppm, I think this was an estimate for the recently filled standard without proper calibration.
20171019 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 460.4 | N/A | std Calibrated 10/17/2017
20171212 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 460.4 | N/A
20180116 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 460.4 | N/A
20180208 | N/A | WG1 | 5/5/2014 | Final pCO2 coefficients A: -30.2789; Final pCO2 coefficients B: 0.43639; Final pCO2 coefficients C: 1.092; Final pCO2 coefficients D: -0.002425 | 460.4 | N/A | Standard postcal 3/14/18 462.03 ppm
20180323 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 496.9 | 2018MAR21 | Updated with new calibrations 3-29. New standard fill also.
20180516 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 496.9 | 2018MAR21
20180710 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 496.9 | 2018MAR21
20180730 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 496.9 | 2018MAR21
20180829 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 474.46 | 2018AUG27 | CO2 stnd refilled and system serviced!
20181001 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 474.46 | 2018AUG27
20181030 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 474.46 | 2018AUG27 | CO2 air intake filters replaced
20181210 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 474.46 | 2018AUG27
20181217 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 474.46 | 2018AUG27
20190114 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 474.46 | 2018AUG27 | STD post cal 476.65 06FEB2019
20190213 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 473.11 | 2019FEB11 | Standard filled refilled on 07FEB2019
20190304 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 473.11 | 2019FEB11
20190502 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 473.11 | 2019FEB11
20190513 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 473.11 | 2019FEB11
20190627 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 473.11 | 2019FEB11 | External clear vent line damaged, replaced!
20190719 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 473.11 | 2019FEB11
20190909 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 473.11 | 2019FEB11
20191001 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 473.11 | 2019FEB11 | Standard gas exhausted… | pCO2 system stopped working on this deployment
20191114 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 503.6 | 2019NOV12 | Standard tank refilled and caled, should be caled again after deployment, std leaked during this deployment. Standard refilled
20191213 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 491.14 | 2019DEC10
20200106 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 491.14 | 2019DEC10
20200203 | N/A | WG1 | 3/29/2018 | Final pCO2 coefficients A: -15.5628; Final pCO2 coefficients B: 0.3292; Final pCO2 coefficients C: 1.0433; Final pCO2 coefficients D: -0.0013 | 491.14 | 2019DEC10
20200227 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 491.14 | 2019DEC10 | The CO2 sensor was calibrated and serviced before this deployment
20200708 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 491.14 | 2019DEC10
20200717 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 491.14 | 2019DEC10
20200903 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 491.14 | 2019DEC10 | 489.74 ppm post call on standard before refill
20201006 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 460.4 | 2020OCT01 | Standard refilled, rough cal, needs post cal
20201202 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 460.4 | 2020OCT01
20210209 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 460.4 | 2020OCT01 | post cal of 461.4ppm 2021-02-25
20210318 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 460.4 | 2020OCT01 | No OA package was deployed, constants just for processing script to run
20210408 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 456.9 | 2021APR01 | Standard swapped, tank WG#1 installed, pCO2 sampling at 30 min intervals for this deployment. We noticed a timing issue with pCO2 data during predeployment testing, check to make sure pCO2 data was at 30 min intervals.
20210528 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 456.9 | 2021APR01 | pCO2 sampling back at 60 min intervals
20210701 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 456.9 | 2021APR01
20210804 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 456.9 | 2021APR01 | Standard tank may be exhausted?
20210818 | N/A | WG1 | 2/20/2020 | Final pCO2 coefficients A: -25.004; Final pCO2 coefficients B: 0.8171; Final pCO2 coefficients C: 1.0657; Final pCO2 coefficients D: -0.0019 | 456.9 | 2021APR01 | No OA package was deployed, copied cells to keep script working, payload is onshore for calibration
20210902 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 447.1 | 2021Aug18 | Licor calibrated, Standard replaced (only 3/4 full due to leak on bench), new licor coeficients not added to spreadsheet until 08SEP. Reprocess pco2 data with updated coefficients. Gortex filters repalced and sytem leak checked before deployment.
20210928 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 447.1 | 2021Aug18
20211028 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 447.1 | 2021Aug18
20211129 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 447.1 | 2021Aug18
20211215 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 447.1 | 2021Aug18
20220107 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 447.1 | 2021Aug18 | STD at about 500 psi
20220124 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 447.1 | 2021Aug18 | STD at about 500 psi (448.1 post cal)
20220310 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 420.6 | 2022Mar9 | Standard almost empty, refilled and paylaod serviced.
20220420 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 420.6 | 2022Mar9
20220516 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 420.6 | 2022Mar9
20220524 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 420.6 | 2022Mar9
20220602 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 420.6 | 2022Mar9 | std post cal :421.3ppm
20220616 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 423.6 | 2022JUN15 | Standard replaced
20220719 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 423.6 | 2022JUN15 | Equil pump on pressure starting to decline, replace filters for next deployment!
20220909 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 423.6 | 2022JUN15
20220919 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 423.6 | 2022JUN15 | Low equilibrator pressure alarms, change filters! Post cal on 2022OCT13 was 424.7 ppm
20221028 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 461.5 | 2022OCT13 | Filters and co2 standard tank repalced
20221103 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 461.5 | 2022OCT13
20221209 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 461.5 | 2022OCT13
20221214 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 461.5 | 2022OCT13
20230118 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 461.5 | 2022OCT13
20230215 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 461.5 | 2022OCT13
20230302 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 461.5 | 2022OCT13
20230406 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 461.5 | 2022OCT13 | pCO2 standard gas tank exhausted during this deployment. Replaced tank when back onshore and serviced co2 instrument.
20230512 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08
20230605 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08
20230710 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08
20230731 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08
20230908 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08
20231013 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08
20231114 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08
20231208 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08
20231220 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08 | pCO2 payload removed for this deployment
20240123 | N/A | WG1 | 8/18/2021 | Final pCO2 coefficients A: -13.7773; Final pCO2 coefficients B: 0.3869; Final pCO2 coefficients C: 1.0335; Final pCO2 coefficients D: -0.0017 | 429.4 | 2023MAY08 | Std at 900 PSI, Licor needs cal, will do after this deployment
20240215 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 429.4 | 2023MAY08 | pCO2 system calibrated. Gortex filters changed and plumbing dried out.
20240311 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 429.4 | 2023MAY08
20240418 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 429.4 | 2023MAY08
20240522 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY | pCO2 standard tank repalced, Soda Lime and Mole Seive replaced.
20240618 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20240719 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20240813 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20240911 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20241024 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20241108 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20241216 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20250107 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20250131 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20250403 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY
20250514 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 440.4 | 2024MAY | Standard exhausted replace for next mission.
20250606 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 429.1 | 2025JUN02 | Standard repalced and minor service (Dririte and soda lime replaced)
20250717 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 429.1 | 2025JUN02
20250813 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 429.1 | 2025JUN02
20250904 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 429.1 | 2025JUN02
20251020 | N/A | WG1 | 2/07/2024 | Final pCO2 coefficients A: -10.6816; Final pCO2 coefficients B: 0.021193; Final pCO2 coefficients C: 1.0739; Final pCO2 coefficients D: -0.0022204 | 429.1 | 2025JUN02
wgSparky
20160523 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed
20160630 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed
20180531 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed
20180821 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed
20180905 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed
20181106 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed
20181113 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed, vehicle Sparky was damged and lost during this deployment
wgHansen
20190415 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20190522 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20190625 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20190812 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20190927 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20191209 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20191218 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20200619 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20200716 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20200909 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20201005 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20201113 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20201204 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20210218 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed. |  | N/A | N/A | No OA package was deployed.
20210319 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20210409 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20210518 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20210625 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20210810 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20211026 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220118 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220202 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220209 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220216 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220323 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220420 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220525 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220705 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220815 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220908 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20220929 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20221020 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20221109 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20221130 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230112 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230209 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230404 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230515 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230607 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230707 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230809 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230901 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20230927 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20231018 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20231031 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20231108 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20231130 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240118 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240228 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240415 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240516 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240708 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240807 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240911 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20241007 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20241119 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20250107 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20250311 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20250415 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20250527 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20250709 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20250820 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20251001 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
wgTest
20230509 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20231129 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240304 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20240321 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.
20250604 | N/A | N/A | N/A | Final pCO2 coefficients A: N/A; Final pCO2 coefficients B: N/A; Final pCO2 coefficients C: N/A; Final pCO2 coefficients D: N/A | N/A | N/A | No OA package was deployed.

## pH
Deployment | Inventory info |  |  |  | Calibration info | [5 empty] | Constants/Comments
| Cap Adapter | Electrode | Model | Housing | Cal date | Batch # | pH of stnd | Vout | Final pH coefficients |  | R | F
wgOA | [8 empty] | Ts | E0 | 8.3145 | 96487
20120316 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | Not sure what the coeff's were for this depl, but GF's processed pH data are at: //atlas/ProjectLibrary/901310_WaveGlider_Feasibility/SV2/SV2-OA/deployment_data/WaveGlider_pH.xlsx
20120521 | N/A | N/A | N/A | N/A | 7/11/2012 | 115 | 7.989 | 0.0471 | 291.15 | -0.41161 | These agree very well with CO2.
20120829 | N/A | N/A | N/A | N/A | 10/2/2012 | 115 | 8.012 | 0.05034 | 293.59 | -0.41639 | These agree better than the 8/30/2012 cal.
20121015 | N/A | N/A | N/A | N/A | 10/2/2012 | 115 | 8.012 | 0.05034 | 293.59 | -0.41639
20130123 | N/A | N/A | N/A | N/A | 1/9/2013 | 115 | 8.012 | 0.07985 | 293.47 | -0.38669 | New electrode
20130315 | N/A | N/A | N/A | N/A | 1/9/2013 | 115 | 8.012 | 0.07985 | 293.47 | -0.38669
20130506 | N/A | N/A | N/A | N/A | 1/9/2013 | 115 | 8.012 | 0.07985 | 293.47 | -0.38669
20130625 | N/A | N/A | N/A | N/A | 6/24/2013 | 126 | 7.999 | 0.0793 | 293.82 | -0.38704
20130909 | N/A | N/A | N/A | N/A | 6/24/2013 | 126 | 7.999 | 0.0793 | 293.82 | -0.38704
20141217 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | This was an early/test deployment of the new WG2 system and it did not collect pH data.
20150825 | N/A | 1237Y200000021846 | N/A | N/A | 6/2/2015 | 136 | 8.2261 | 0.08688 | 279.0284 | -0.36855 | Extra cal of same sensor as 20150923
20150923 | N/A | N/A | N/A | N/A | 9/15/2015 | 144 | 8.1878 | 0.08327 | 279.1379 | -0.37022 | First deployment with 'upgraded' pH/pCO2 system.
wgTiny
20140318 | N/A | N/A | N/A | N/A | 9/10/2013 | 126 | 7.991 | 0.0729 | 294.35 | -0.39381 | NO PH DATA ON FETCH. MAY BE RAW PH DATA ON CF.
20140403 | N/A | N/A | N/A | N/A | 9/10/2013 | 126 | 7.991 | 0.0729 | 294.35 | -0.39381 | NO PH DATA ON FETCH. RAW PH DATA ON CF ONLY.
20140521 | N/A | N/A | N/A | N/A | 9/10/2013 | 126 | 7.991 | 0.0729 | 294.35 | -0.39381 | pH V1 is really low ~1 mV (an order of mag low) and appears random/with no clear signal.
20140619 | N/A | N/A | N/A | N/A | 9/10/2013 | 126 | 7.991 | 0.0729 | 294.35 | -0.39381 | pH V1 is really low ~1 mV (an order of mag low) and appears random/with no clear signal.
20140918 | N/A | N/A | N/A | N/A | 9/10/2013 | 126 | 7.991 | 0.0729 | 294.35 | -0.39381 | Although 8/21/14 is a more recent cal, this gives pH ~8.2; whereas 9/10/13 cal gives pH ~8.0. Estimated pCO2 (from derived pH) is still off by ~50 uatm.
20150212 | N/A | N/A | N/A | N/A | 9/10/2013 | 126 | 7.991 | 0.0729 | 294.35 | -0.39381 | Although 8/21/14 is a more recent cal, when compared to (1) M1 (when nearby) and (2) estimated pH from pCO2water, the 9/10/2013 cal gives better results.
20150421 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | New electrode. Used an opened batch from Virginia Elrod. Other standards were in G of C. No usable pH data this deployment, pump issues.
20150501 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979
20150611 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | Payload 3 (OA) went out after first few days…
20150824 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | pH data are ~0.05-0.1 unit high compared to predicted from pCO2.
20150921 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979
20151019 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | No OA data due to CTD payload failure
20151209 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | Converted pH data are way too high, 0.1 higher than M1, and don't agree with pCO2w.
20160128 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | pH voltages are all < 0.025 V, possibly dead batteries. No CTD data, so pCO2 temperatures used in derivation of pCO2w and pH.
20160225 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | No science data collected this deployment. Apparently a new pH electrode was swapped in for this deployment, 1508Y200000063189, calibrated 3/7/2016.
20160315 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | pH vs pCO2w plot does not look good; either pH or pCO2 calibration is off. **Apparently used a different pH electrode, same as prev depl, *3189.
20160627 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20160701 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20160830 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20160920 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20161005 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20161031 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20161212 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20170109 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20170307 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20170412 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20170605 | N/A | N/A | N/A | N/A | 4/20/2015 | 136 | 7.9316 | 0.06341 | 293.12 | -0.3979 | **Apparently used a different pH electrode, same as prev depl, *3189.
20170630 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 6/26/2017 | 163 | 7.8877 | 0.07075 | 294.01 | -0.3894 | Swapped this electrode in even though previous deployments data looked good.
20170823 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 6/26/2017 | 163 | 7.8877 | 0.07075 | 294.01 | -0.3894
20170921 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 6/26/2017 | 163 | 7.8877 | 0.07075 | 294.01 | -0.3894
20171019 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 6/26/2017 | 163 | 7.8877 | 0.07075 | 294.01 | -0.3894
20171212 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 6/26/2017 | 163 | 7.8877 | 0.07075 | 294.01 | -0.3894
20180116 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 6/26/2017 | 163 | 7.8877 | 0.07075 | 294.01 | -0.3894
20180208 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 6/26/2017 | 163 | 7.8877 | 0.07075 | 294.01 | -0.3894 | pH had isssues on 03/06/2018
20180323 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871 | pH was calibrated before deployment, coefficients updated on 4/2 readings off TA lines? Used calibration form 5/14/18 calibration, old calibration is pasted to right |  | 3/22/2018 | 171 | 7.9012 | 0.06563 | 295.2053 | -0.39718
20180516 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20180710 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871 | pH plumbing issues, data bad
20180730 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20180829 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20181001 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20181030 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20181210 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20181217 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20190114 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20190213 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20190304 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20190502 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871
20190513 | N/A | 1633Y200000076201 | 51205554-501 | N/A | 5/14/2018 | 173 | 7.8945 | 0.07126 | 292.617 | -0.3871 | pH data bad intermittently. It looks like the electrode failed. It has been replaced
20190627 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986 | pH sensor and cap adapter were replaced
20190719 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986
20190909 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986 | pH issues and GPCTD issues, bad pump?
20191001 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986 | GPCTD pump replaced, there was a cable failure and GPCTD cable was replaced, pH might be out of batteries?
20191114 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986 | pH batteries replaced
20191213 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986
20200106 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986
20200203 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986
20200227 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986 | pH data looks bad…
20200708 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986
20200717 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986
20200903 | N/A | 1807Y200000082424 | 51205554-501 | N/A | 6/26/2019 | 179 | 8.0972 | 0.0729 | 293.47 | -0.3986 | bad pH?
20201006 | N/A | 1808Y200000083055 | 51205554-501 | N/A | 10/1/2020 | 187 | 7.9634 | 0.06366 | 292.86 | -0.39908 | new pH electrode and batteries. Calibrated
20201202 | N/A | 1808Y200000083055 | 51205554-501 | N/A | 10/1/2020 | 187 | 7.9634 | 0.06366 | 292.86 | -0.39908
20210209 | N/A | 1808Y200000083055 | 51205554-501 | N/A | 10/1/2020 | 187 | 7.9634 | 0.06366 | 292.86 | -0.39908 | noisy ph data after about 48 hours, looks like a bad wire/solder joint
20210318 | N/A | 1808Y200000083055 | 51205554-501 | N/A | 10/1/2020 | 187 | 7.9634 | 0.06366 | 292.86 | -0.39908 | No OA package was deployed
20210408 | N/A | 1808Y200000083055 | 51205554-501 | N/A | 10/1/2020 | 187 | 7.9634 | 0.06366 | 292.86 | -0.39908 | pH repaired and batteries replaced! pH sampling at 5 min intervals for this deployment, check dataset to make sure pH data is at 5 min intervals
20210528 | N/A | 1808Y200000083055 | 51205554-501 | N/A | 5/26/2021 | 133 | 7.0876 | 0.04326 | 292.63 | -0.36827 | pH cleaned, Yui says that there is a slight offset, shall we recalibrate? pH sampling back at 10 min intervals for this deployment, pH reading look off after calibration! Also pH sensor failed during this mission…
20210701 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874 | pH sensor issues during the last deployment, ph eectrode and cap adapter swapped! Electrode was DOA. Second electrode, original cap adapter, calibrated with tris buffer from Joseph Warren, plumbing clogged temporaily on 06JUL
20210804 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874
20210818 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874 | No OA package was deployed, copied cells to keep script working
20210902 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874 | pH batteries replaced
20210928 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874 | pH sensor face cleaned, GPCTD pump replaced, air bleed hole added to pH manifold.
20211028 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874
20211129 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874 | pH failed on this deployment… Data from the other sensors looks good!
20211215 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874 | pH sensor checked by Jules before this deployment. Data still doesn't look good…
20220107 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874 | One of the pH sensor batteries died. Replaced. Calble and sensor checked.
20220124 | N/A | 1213Y200000092681 | 51205554-501 | N/A | 6/30/2021 | TRIS#6 | 8.2568 | 0.0715 | 293.13 | -0.40874 | pH sensor electrode cleaned before deployment.
20220310 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006 | pH sensor replaced and calibrated, pH sensor reading low during this deployment
20220420 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006 | ph sensor face cleaned for this deployment
20220516 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006 | ph sensor face cleaned for this deployment
20220524 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20220602 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20220616 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006 | pH batteries replaced and electrode face cleaned
20220719 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006 | ph electrode cleaned before this deployment
20220909 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20220919 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20221028 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006 | pH batteries replaced since the co2 payload was opened for work.
20221103 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20221209 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20221214 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20230118 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20230215 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20230302 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006
20230406 | N/A | 2146Y200000094976 |  |  | 3/9/2022 | TRIS#4 | 7.44 | 0.03068 | 292.15 | -0.4006 | pH batteries replaced since the co2 payload was open for work. pH Sensor face also cleaned.
20230512 | N/A | 2146Y200000094976 |  |  | 5/11/2023 | TRIS#4 | 7.3793 | 0.027 | 292.1399 | -0.40075 | pH sensor recalibrated
20230605 | N/A | 2146Y200000094976 |  |  | 5/11/2023 | TRIS#4 | 7.3793 | 0.027 | 292.1399 | -0.40075
20230710 | N/A | 2146Y200000094976 |  |  | 5/11/2023 | TRIS#4 | 7.3793 | 0.027 | 292.1399 | -0.40075
20230731 | N/A | 2146Y200000094976 |  |  | 5/11/2023 | TRIS#4 | 7.3793 | 0.027 | 292.1399 | -0.40075
20230908 | N/A | 2146Y200000094976 |  |  | 5/11/2023 | TRIS#4 | 7.3793 | 0.027 | 292.1399 | -0.40075 | pH sensor failed
20231013 | N/A | 2146Y200000094976 |  |  | 10/11/2023 | TRIS#4 | 7.3767 | 0.0414 | 292.0602 | -0.38608 | pH sensor failed on last deployment, gell refill
20231114 | N/A | 2146Y200000094976 |  |  | 10/11/2023 | TRIS#4 | 7.3767 | 0.0414 | 292.0602 | -0.38608 | pH sensor failed again on this deployment, voltage on one of the battries was low (3.0 Volts)
20231208 | N/A | 2146Y200000094976 |  |  | 10/11/2023 | TRIS#4 | 7.3767 | 0.0414 | 292.0602 | -0.38608 | pH batteries replaced. pH sensors confirmed operational by Jules. pH sensor not working again on this deployment.
20231220 | N/A | 2146Y200000094976 |  |  | 10/11/2023 | TRIS#4 | 7.3767 | 0.0414 | 292.0602 | -0.38608 | pH sensor removed for this deployment.
20240123 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057 | troubleshooting ph issues, durafet is from Jupiter #3, new cap adapter too, pH batteries replaced, A/D confirmed with calibrated voltage source
20240215 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20240311 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20240418 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20240522 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20240618 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20240719 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20240813 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057 | pH sensor reporting bad values, troubleshoot
20240911 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057 | pH batteries were dead, repalced batteries and Jules confirmed that pH sensor itself is functional.
20241024 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20241108 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20241216 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20250107 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20250131 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20250403 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20250514 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20250606 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20250717 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20250813 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20250904 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
20251020 | N/A | 1911Y200000087885 |  |  | 1/22/2024 | TRIS#4 | 7.4016 | 0.0295 | 292.84 | -0.40057
wgSparky
20160523 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed
20160630 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed
20180531 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed
20180821 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed
20180905 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed
20181106 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed
20181113 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed, vehicle Sparky was damged and lost during this deployment
wgHansen
20190415 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20190522 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20190625 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20190812 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20190927 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20191209 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20191218 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20200619 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20200716 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20200909 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20201005 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20201113 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20201204 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20210218 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20210319 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20210409 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20210518 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20210625 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20210810 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20211026 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220118 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220202 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220209 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220216 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220323 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220420 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220525 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220705 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220815 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220908 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20220929 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20221020 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20221109 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20221130 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230112 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230209 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230404 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230515 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230607 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230707 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230809 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230901 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20230927 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20231018 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20231031 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20231108 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20231130 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240118 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240228 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240415 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240516 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240708 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240807 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240911 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20241007 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20241119 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20250107 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20250311 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20250415 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20250527 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20250709 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20250820 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20251001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
wgTest
20230509 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20231129 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240304 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20240321 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.
20250604 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No OA package was deployed.

## BB
Deployment | Inventory info |  |  | Calibration info | [6 empty] | Comments
| BB2FB or Eco-Puck S/N | Lambda1 | Lambda2 | Cal date | Cal coefficients
wgOA |  |  |  |  | Lambda1 SF | Lambda1 Dark | Lambda2 SF | Lambda2 Dark | Chl SF | Chl Dark
20120316 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO BBF2B DATA
20120521 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO BBF2B DATA
20120829 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO BBF2B DATA
20121015 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO BBF2B DATA
20130123 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO BBF2B DATA
20130315 | BBF2B-227 | 470 | 700 | 6/9/2014 | 3.64e-05 | 70 | 3.874e-06 | 44 | 0.0151 | 55 | BBF2B data logged internally. 'bb' values are all negative, 'beta' are all positive
20130506 | BBF2B-227 | 470 | 700 | 6/9/2014 | 3.64e-05 | 70 | 3.874e-06 | 44 | 0.0151 | 55 | BBF2B data logged internally. 'bb' values are all negative, 'beta' are all positive
20130625 | BBF2B-227 | 470 | 700 | 6/9/2014 | 3.64e-05 | 70 | 3.874e-06 | 44 | 0.0151 | 55 | BBF2B data logged internally. 'bb' values are all negative, 'beta' are all positive
20130909 | BBF2B-227 | 470 | 700 | 6/9/2014 | 3.64e-05 | 70 | 3.874e-06 | 44 | 0.0151 | 55 | BBF2B data logged internally. 'bb' values are all negative, 'beta' are all positive
20141217 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO BBF2B DATA
20150825 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO BBF2B DATA (dock test only)
20150923 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | BBF2B data logged internally?
wgTiny
20140318 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO ECO-PUCK DATA ON FETCH
20140403 | BB2FB-227 | 470 | 700 | 6/9/2014 | 3.64e-05 | 70 | 3.874e-06 | 44 | 0.0151 | 55 | NO ECO-PUCK DATA ON FETCH
20140521 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO BB2FB DATA -- OUT FOR CALIBRATION
20140619 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | NO ECO-PUCK DATA ON FETCH
20140918 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49
20150212 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49 | NO ECO-PUCK DATA -- BOARD FAILURE
20150421 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49
20150501 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49
20150611 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49
20150824 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49
20150921 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49
20151019 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49
20151209 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49
20160128 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No BB data collected this deployment
20160225 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No science data collected this deployment
20160315 | Eco-Puck BB2FLMBA-1133 | 470 | 650 | 12/4/2013 | 1.08e-05 | 47 | 3.949e-06 | 46 | 0.0121 | 49 | A new Eco-Puck will be used for the next wgTiny mission.
20160627 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35 | New Eco-Puck with copper anti-foul faceplate. On the vehicle hull, not the glider.
20160701 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20160830 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20160920 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20161005 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20161031 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20161212 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20170109 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20170307 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20170412 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20170605 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 4/21/2016 | 1.098e-05 | 52 | 3.629e-06 | 40 | 0.012 | 35
20170630 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 1/24/2017 | 1.062e-05 | 50 | 3.572e-06 | 45 | 0.0121 | 44 | Swapped-in for this depl.
20170823 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 1/24/2017 | 1.062e-05 | 50 | 3.572e-06 | 45 | 0.0121 | 44
20170921 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 1/24/2017 | 1.062e-05 | 50 | 3.572e-06 | 45 | 0.0121 | 44 | Eco Puck started showing signs of cable failure
20171019 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 1/24/2017 | 1.062e-05 | 50 | 3.572e-06 | 45 | 0.0121 | 44 | Eco Puck stopped working on this deployment, appears to be from power outtage
20171212 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 1/24/2017 | 1.062e-05 | 50 | 3.572e-06 | 45 | 0.0121 | 44 | Did not have time to troubleshoot Ecopuck
20180116 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 7/17/2017 | 1.046e-05 | 54 | 4.847e-06 | 48 | 0.0119 | 39 | Cable failure on 20171019 deployment, replaced Ecopuck just in case
20180208 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 7/17/2017 | 1.046e-05 | 54 | 4.847e-06 | 48 | 0.0119 | 39 | vehicle fouled with fishing line, seaweed and floats, FL/BB data might be showing signs of this starting on 07MAR2018?
20180323 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 7/17/2017 | 1.046e-05 | 54 | 4.847e-06 | 48 | 0.0119 | 39
20180516 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 7/17/2017 | 1.046e-05 | 54 | 4.847e-06 | 48 | 0.0119 | 39 | ECO/VR2C cable failed during testing replaced
20180710 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 7/17/2017 | 1.046e-05 | 54 | 4.847e-06 | 48 | 0.0119 | 39
20180730 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47 | Cable failure on last deployment? Whats going on? Cable and ecopuck replaced
20180829 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20181001 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20181030 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20181210 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20181217 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20190114 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20190213 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20190304 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20190502 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20190513 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47
20190627 | Eco-Puck BB2FLMBC-1541 | 470 | 650 | 2/14/2018 | 1.135e-05 | 60 | 4.3e-06 | 53 | 0.0126 | 47 | Cable failure EcoPcuk not working
20190719 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38 | Sensor and cable replaced!
20190909 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20191001 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38 | EcoPuck moved forward in wet bay, to accommodate MF transducer
20191114 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20191213 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20200106 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20200203 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20200227 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20200708 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20200717 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20200903 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20201006 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20201202 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20210209 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20210318 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20210408 | Eco-Puck BB2FLMBC-1429 | 470 | 650 | 10/25/2018 | 1.088e-05 | 49 | 5.927e-06 | 45 | 0.0113 | 38
20210528 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49 | Brand new Ecopuck installed due to old calibration on 1429
20210701 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20210804 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20210818 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20210902 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20210928 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20211028 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20211129 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49 | Diffuculties starting ecopuck during this mission, cable issues?
20211215 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20220107 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20220124 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49 | Lots of barnacles around Ecopuck and the data looked noisy towards end of deployment
20220310 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49 | Ecopuck cleaned and copper foil tape added, sensor stopped working on this deployment
20220420 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49 | Cable replaced.
20220516 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20220524 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20220602 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49
20220602 | Eco-Puck BB2FLMBC-5678 | 470 | 650 | 5/14/2019 | 1.144e-05 | 39 | 4.098e-06 | 39 | 0.012 | 49 | connector cleaned and relubed
20220719 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39 | Ecopuck swapped with brand new sensor, cal info updated!
20220909 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39 | connector serviced while installing JRF hull
20220919 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20221028 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20221103 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20221209 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20221214 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20230118 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20230215 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20230302 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20230406 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39 | Ecopuck connector serviced while working on co2 payload.
20230512 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20230605 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20230710 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20230731 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20230908 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20231013 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20231114 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20231208 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20231220 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20240123 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 6/30/2022 | 1.131e-05 | 42 | 3.596e-06 | 49 | 0.012 | 39
20240215 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50 | Sensor 5949 installed on 05FEB2024, working but timestamps bad?
20240311 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50 | Memory full on Science Payload SMC. Removed some logs and everything is working now!
20240418 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20240522 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20240618 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20240719 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20240813 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50 | Ecopuck connector regreased.
20240911 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20241024 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20241108 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20241216 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20250107 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 7/31/2023 | 1.133e-05 | 53 | 3.683e-06 | 51 | 0.0121 | 50
20250131 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 3/5/2024 | 1.145e-05 | 40 | 4.261e-06 | 48 | 0.0121 | 39 | Installed freshly calibrated Ecopuck
20250403 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 3/5/2024 | 1.145e-05 | 40 | 4.261e-06 | 48 | 0.0121 | 39
20250514 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 3/5/2024 | 1.145e-05 | 40 | 4.261e-06 | 48 | 0.0121 | 39
20250606 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 3/5/2024 | 1.145e-05 | 40 | 4.261e-06 | 48 | 0.0121 | 39
20250717 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 3/5/2024 | 1.145e-05 | 40 | 4.261e-06 | 48 | 0.0121 | 39
20250813 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 3/5/2024 | 1.145e-05 | 40 | 4.261e-06 | 48 | 0.0121 | 39
20250904 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 3/5/2024 | 1.145e-05 | 40 | 4.261e-06 | 48 | 0.0121 | 39
20251020 | Eco-Puck BB2FLMBC-7618 | 470 | 650 | 3/5/2024 | 1.145e-05 | 40 | 4.261e-06 | 48 | 0.0121 | 39
wgSparky
20160523 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No BB sensor was deployed
20160630 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No BB sensor was deployed
20180531 | Eco-Puck BB2FLMBC-1648 | 470 | 650 | 12/26/2017 | 1.17e-05 | 55 | 4.104e-06 | 39 | 0.0121 | 46 | Installed this new sensor in May 2018, confirm!
20180821 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No BB sensor was deployed
20180905 | Eco-Puck BB2FLMBC-1648 | 470 | 650 | 12/26/2017 | 1.17e-05 | 55 | 4.104e-06 | 39 | 0.0121 | 46 | Installed this new sensor in May 2018, confirm!
20181106 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No BB sensor was deployed
20181113 | Eco-Puck BB2FLMBC-1648 | 470 | 650 | 12/26/2017 | 1.17e-05 | 55 | 4.104e-06 | 39 | 0.0121 | 46 | Reinstalled sensor, vehicle Sparky was damged and lost during this deployment
wgHansen
20190415 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20190522 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20190625 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20190812 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20190927 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20191209 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20191218 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20200619 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20200716 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20200909 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20201005 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20201113 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20201204 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20210218 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20210319 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20210409 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20210518 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20210625 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20210810 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20211026 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220118 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220202 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220209 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220216 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220323 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220420 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220525 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220705 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220815 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220908 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20220929 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20221020 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20221109 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20221130 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230112 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230209 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230404 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230515 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230607 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230707 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230809 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230901 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20230927 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20231018 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20231031 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20231108 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20231130 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20240118 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20240228 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20240415 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20240516 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20240708 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20240807 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20240911 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20241007 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20241119 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20250107 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20250311 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20250415 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20250527 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20250709 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20250820 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
20251001 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | No EcoPuck deployed
wgTest
20230509 | [10 empty] | Ecopuck installed for test only, don't think it was running though?
20231129 | [10 empty] | Ecopuck installed for test only, don't think it was running though?
20240304 | [10 empty] | Ecopuck installed for test only, don't think it was running though?
20240321 | [10 empty] | Ecopuck installed for test only, don't think it was running though?
20250604 | Eco-Puck BB2FLMBC-5949 | 470 | 650 | 3/20/2025 | 1.3e-05 | 48 | 5.296e-06 | 42 | 0.0121 | 43 | SN5949 installed.

## VR2C
Deployment | Inventory info | Comments
Receiver #
wgOA
20120316 | N/A
20120521 | N/A
20120829 | N/A
20121015 | N/A
20130123 | N/A
20130315 | N/A
20130506 | N/A
20130625 | N/A
20130909 | N/A
20141217 | N/A
20150825 | N/A
20150923 | N/A | First wgOA deployment with a VR2C. VR2C was removed from wgOA after this mission to install on Hopkins (OA1) turn-around, summer 2016.
wgTiny
20140318 | N/A
20140403 | N/A
20140521 | N/A
20140619 | N/A
20140918 | N/A
20150212 | N/A
20150421 | N/A
20150501 | N/A
20150611 | N/A
20150824 | N/A
20150921 | N/A | First SV3 deployment with a VR2C.
20151019 | N/A
20151209 | N/A
20160128 | N/A | No VR2C data collected this deployment
20160225 | N/A | No science data collected this deployment
20160315 | N/A
20160627 | 450157 | A test tag (ID = 10, 11, 12) was tested successfully on shore prior to deployment. Time zone for tag hits may be in PST, whereas status messages appear to be GMT. No hits this deployment.
20160701 | 450157
20160830 | 450157
20160920 | 450157
20161005 | 450157
20161031 | 450157 | Three hits while near shipping lane off Big Sur, but likely false detections
20161212 | 450157
20170109 | 450157 | One hit near Station C1, but likely a false detection
20170307 | 450157
20170412 | 450157 | Several hits, but all were confirmed as false.
20170605 | 450157
20170630 | 450157
20170823 | 450157
20170921 | 450157
20171019 | 450157
20171212 | 450157
20180116 | 450157
20180208 | 450157
20180323 | 450157
20180516 | 450157 | One false detect and then lots of testing for Hotpsot/Jelly fish tracking project
20180710 | 450157
20180730 | 450157 | Received new test tags!
20180829 | 450157
20181001 | 450157
20181030 | 450157
20181210 | 450157
20181217 | 450157
20190114 | 450157
20190213 | 450157
20190304 | 450157
20190502 | 450157
20190513 | 450157
20190627 | 450157
20190719 | 450157
20190909 | 450157
20191001 | 450157
20191114 | 450157
20191213 | 450157
20200106 | 450157
20200203 | 450157
20200227 | 450157
20200708 | 450157 | Tag 15877 (test tag) was left on Tiny for this deployment. Looks like we got a hit from tag 56500 also?
20200717 | 450157
20200903 | 450157
20201006 | 450157
20201202 | 450157
20210209 | N/A | no VR2C was deployed
20210318 | 450157
20210408 | 450157
20210528 | 450157
20210701 | 450157
20210804 | 450157 | Internalcoin battery dead?
20210818 | 450157
20210902 | 450157
20210928 | 450157
20211028 | 450157
20211129 | 450157
20211215 | 450157
20220107 | 450157
20220124 | 450157
20220310 | 450157 | VR2C Removed for this deployment
20220420 | 450157 | VR2C Removed for this deployment
20220516 | 450157 | VR2C Removed for this deployment
20220524 | 450157 | VR2C Removed for this deployment
20220602 | 450157 | VR2C Removed for this deployment
20220616 | 450157 | VR2C Removed for this deployment
20220719 | 450157 | VR2C Removed for this deployment
20220909 | 450157 | VR2C Removed for this deployment
20220919 | 450157 | VR2C Removed for this deployment
20221028 | 450157 | VR2C Removed for this deployment
20221103 | 450157 | VR2C Removed for this deployment
20221209 | 450157 | VR2C Removed for this deployment
20221214 | 450157 | VR2C Removed for this deployment
20230118 | 450157 | VR2C Removed for this deployment
20230215 | 450157 | VR2C Removed for this deployment
20230302 | 450157 | VR2C Removed for this deployment
20230406 | 450157 | VR2C Removed for this deployment
20230512 | 450157 | VR2C Removed for this deployment
20230605 | 450157 | VR2C Removed for this deployment
20230710 | 450157 | VR2C Removed for this deployment
20230731 | 450157 | VR2C Removed for this deployment
20230908 | 450157 | VR2C Removed for this deployment
20231013 | 450157 | VR2C Removed for this deployment
20231114 | 450157 | VR2C Removed for this deployment
20231208 | 450157 | VR2C Removed for this deployment
20231220 | 450157 | VR2C Removed for this deployment
20240123 | 450157 | VR2C Removed for this deployment
20240215 | 450157 | VR2C Removed for this deployment
20240311 | 450157 | VR2C Removed for this deployment
20240418 | 450157 | VR2C Removed for this deployment
20240522 | 450157 | VR2C Removed for this deployment
20240618 | 450157 | VR2C Removed for this deployment
20240719 | 450157 | VR2C Removed for this deployment
20240813 | 450157 | VR2C Removed for this deployment
20240911 | 450157 | VR2C Removed for this deployment
20241024 | 450157 | VR2C Removed for this deployment
20241108 | 450157 | VR2C Removed for this deployment
20241216 | 450157 | VR2C Removed for this deployment
20250107 | 450157 | VR2C Removed for this deployment
20250131 | 450157 | VR2C Removed for this deployment
20250403 | 450157 | VR2C Removed for this deployment
20250514 | 450157 | VR2C Removed for this deployment
20250606 | 450157 | VR2C Removed for this deployment
20250717 | 450157 | VR2C Removed for this deployment
20250813 | 450157 | VR2C Removed for this deployment
20250904 | 450157 | VR2C Removed for this deployment
20251020 | 450157 | VR2C Removed for this deployment
wgSparky
20160523 | N/A | No VR2C was deployed
20160630 | N/A | No VR2C was deployed
20180531 | N/A | No VR2C was deployed
20180821 | N/A | No VR2C was deployed
20180905 | N/A | No VR2C was deployed
20181106 | N/A | No VR2C was deployed
20181113 | N/A | No VR2C was deployed, vehicle Sparky was damged and lost during this deployment
wgHansen
20190415 | N/A | No VR2C deployed
20190522 | N/A | No VR2C deployed
20190625 | N/A | No VR2C deployed
20190812 | N/A | No VR2C deployed
20190927 | N/A | No VR2C deployed
20191209 | N/A | No VR2C deployed
20191218 | N/A | No VR2C deployed
20200619 | N/A | No VR2C deployed
20200716 | N/A | No VR2C deployed
20200909 | N/A | No VR2C deployed
20201005 | N/A | No VR2C deployed
20201113 | N/A | No VR2C deployed
20201204 | N/A | No VR2C deployed
20210218 | N/A | No VR2C deployed
20210319 | N/A | No VR2C deployed
20210409 | N/A | No VR2C deployed
20210518 | N/A | No VR2C deployed
20210625 | N/A | No VR2C deployed
20210810 | N/A | No VR2C deployed
20211026 | N/A | No VR2C deployed
20220118 | N/A | No VR2C deployed
20220202 | N/A | No VR2C deployed
20220209 | N/A | No VR2C deployed
20220216 | N/A | No VR2C deployed
20220323 | N/A | No VR2C deployed
20220420 | N/A | No VR2C deployed
20220525 | N/A | No VR2C deployed
20220705 | N/A | No VR2C deployed
20220815 | N/A | No VR2C deployed
20220908 | N/A | No VR2C deployed
20220929 | N/A | No VR2C deployed
20221020 | N/A | No VR2C deployed
20221109 | N/A | No VR2C deployed
20221130 | N/A | No VR2C deployed
20230112 | N/A | No VR2C deployed
20230209 | N/A | No VR2C deployed
20230404 | N/A | No VR2C deployed
20230515 | N/A | No VR2C deployed
20230607 | N/A | No VR2C deployed
20230707 | N/A | No VR2C deployed
20230809 | N/A | No VR2C deployed
20230901 | N/A | No VR2C deployed
20230927 | N/A | No VR2C deployed
20231018 | N/A | No VR2C deployed
20231031 | N/A | No VR2C deployed
20231108 | N/A | No VR2C deployed
20231130 | N/A | No VR2C deployed
20240118 | N/A | No VR2C deployed
20240228 | N/A | No VR2C deployed
20240415 | N/A | No VR2C deployed
20240516 | N/A | No VR2C deployed
20240708 | N/A | No VR2C deployed
20240807 | N/A | No VR2C deployed
20240911 | N/A | No VR2C deployed
20241007 | N/A | No VR2C deployed
20241119 | N/A | No VR2C deployed
20250107 | N/A | No VR2C deployed
20250311 | N/A | No VR2C deployed
20250415 | N/A | No VR2C deployed
20250527 | N/A | No VR2C deployed
20250709 | N/A | No VR2C deployed
20250820 | N/A | No VR2C deployed
20251001 | N/A | No VR2C deployed
wgTest
20230509 | N/A | No VR2C deployed
20231129 | N/A | No VR2C deployed
20240304 | N/A | No VR2C deployed
20240321 | N/A | No VR2C deployed
20250604 | N/A | No VR2C deployed

## GPCTD+DO
Deployment | Float CTD |  |  | Float DO calibration (factory) | [7 empty] | Sub CTD |  | Sub DO calibration (factory) | [7 empty] | Comments
| GPCTD | DO | Pump | Cal date | Coefficients | [6 empty] | GPCTD | DO | Cal date | Coefficients
wgOA |  |  |  |  | Soc | Foffset | Tau20 | A | B | C | E nominal |  |  |  | Soc | Foffset | Tau20 | A | B | C | E nominal
20120316 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | Have not found cal info for wgOA yet, but this was never needed because fetch outputted O2 already in ml/l.
20120521 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20120829 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20121015 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20130123 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20130315 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20130506 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20130625 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20130909 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20141217 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
20150825 | N/A | N/A | N/A | 3/31/2015 | 0.00026506 | -894.46 | 1.18 | -0.0037981 | 0.0001857 | -2.1783e-06 | 0.036 | [10 empty] | Original DO sensor for wgOA. Calibrated 3/6/2015 and re-installed on wgOA.
20150923 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A
wgTiny
20140318 | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | N/A | [10 empty] | A new CTD was purchased for use with wgTiny.
20140403 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036 | [10 empty] | These first few mission dates precede the CTD calibration date??
20140521 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20140619 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20140918 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20150212 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20150421 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036 | [10 empty] | O2 data bad, issues with pump
20150501 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20150611 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20150824 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20150921 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20151019 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20151209 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036
20160128 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036 | [10 empty] | No CTD data collected this deployment
20160225 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036 | [10 empty] | No science data collected this deployment
20160315 | 0118 | 0212 | N/A | 5/21/2014 | 0.00029681 | -787.66 | 1.56 | -0.0027123 | 0.00020315 | -3.8309e-06 | 0.036 | [10 empty] | A new CTD will be used for the next wgTiny mission.
20160627 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036 | [10 empty] | New CTD, fresh factory-cal. CW installed 6/1/2016.
20160701 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20160830 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20160920 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20161005 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20161031 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20161212 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20170109 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20170307 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20170412 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20170605 | 0164 | 3233 | N/A | 9/25/2015 | 0.00031556 | -817.4 | 1.41 | -0.0038249 | 0.00015244 | -2.1156e-06 | 0.036
20170630 | 0118 | 0212 | 5M 051537 | 8/24/2016 | 0.00028881 | -808.26 | 0.91 | -0.0042784 | 0.00021726 | -3.1113e-06 | 0.036
20170823 | 0118 | 0212 | 5M 051537 | 8/24/2016 | 0.00028881 | -808.26 | 0.91 | -0.0042784 | 0.00021726 | -3.1113e-06 | 0.036
20170921 | 0118 | 0212 | 5M 051537 | 8/24/2016 | 0.00028881 | -808.26 | 0.91 | -0.0042784 | 0.00021726 | -3.1113e-06 | 0.036
20171019 | 0118 | 0212 | 5M 051537 | 8/24/2016 | 0.00028881 | -808.26 | 0.91 | -0.0042784 | 0.00021726 | -3.1113e-06 | 0.036
20171212 | 0118 | 0212 | 5M 051537 | 8/24/2016 | 0.00028881 | -808.26 | 0.91 | -0.0042784 | 0.00021726 | -3.1113e-06 | 0.036
20180116 | 0118 | 0212 | 5M 051537 | 8/24/2016 | 0.00028881 | -808.26 | 0.91 | -0.0042784 | 0.00021726 | -3.1113e-06 | 0.036
20180208 | 0118 | 0212 | 5M 051537 | 8/24/2016 | 0.00028881 | -808.26 | 0.91 | -0.0042784 | 0.00021726 | -3.1113e-06 | 0.036 | [10 empty] | Salinity went bad on 2/23/2018, and again on 3/6/2018, along with DO and pH
20180323 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036 | [10 empty] | Replaced ctd with recently calibrated one. Salinity measurements started off normal, but then dropped on 26MAR2018 and recovereed on 29MAR2018
20180516 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20180710 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20180730 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20180829 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20181001 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20181030 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20181210 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20181217 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20190114 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20190213 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20190304 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20190502 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20190513 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20190627 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036
20190719 | 0164 | 03233 | 5M-8405 | 7/12/2017 | 0.00032259 | -826.99 | 0.92 | -0.0044998 | 0.00018691 | -2.802e-06 | 0.036 | [10 empty] | Upon removal of this GPCTD, I noticed that the pump with unit 0164 failed at some point!
20190909 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036 | [10 empty] | Installed a new GPCTD plus DO for this deployment, Cond/Sal measurements low…
20191001 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036 | [10 empty] | Main GPCTD cable failed and was replaced, pump also replaced. GPCTD stopped working on this deployment and then started working again!
20191114 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20191213 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20200106 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20200203 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20200227 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20200708 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20200717 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20200903 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20201006 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20201202 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20210209 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20210318 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20210408 | 0286 | 3803 | 5M-9624 | 3/24/2019 | 0.00028974 | -817.39 | 2 | -0.0034502 | 0.00015233 | -2.3434e-06 | 0.036
20210528 | 0164 | 3233 | 5M-8405 | 10/16/2019 | 0.00032545 | -816.09 | 1.11 | -0.004217 | 0.00018041 | -2.8371e-06 | 0.036 | [10 empty] | GPCTD replaced due to old calibration on the other previous one!
20210701 | 0164 | 3233 | 5M-8405 | 10/16/2019 | 0.00032545 | -816.09 | 1.11 | -0.004217 | 0.00018041 | -2.8371e-06 | 0.036 | [10 empty] | Plumbing clogged temporarily on 06JUL
20210804 | 0164 | 3233 | 5M-8405 | 10/16/2019 | 0.00032545 | -816.09 | 1.11 | -0.004217 | 0.00018041 | -2.8371e-06 | 0.036
20210818 | 0164 | 3233 | 5M-8405 | 10/16/2019 | 0.00032545 | -816.09 | 1.11 | -0.004217 | 0.00018041 | -2.8371e-06 | 0.036 | [10 empty] | GPCTD physically installed for deployment, but not connected with a cable and not retuning data
20210902 | 0164 | 3233 | 5M-8405 | 10/16/2019 | 0.00032545 | -816.09 | 1.11 | -0.004217 | 0.00018041 | -2.8371e-06 | 0.036 | [10 empty] | DO stopped reporting troubleshoot after recovery.
20210928 | 0118 | 0212 | 5M-9624 | 11/30/2020 | 0.00027441 | -795.63 | 1.09 | -0.0045878 | 0.00021026 | -3.3257e-06 | 0.036 | [10 empty] | Pump performance on last deployment was poor and DO failed so GPCTD was swapped for this deployment.
20211028 | 0118 | 0212 | 5M-9624 | 11/30/2020 | 0.00027441 | -795.63 | 1.09 | -0.0045878 | 0.00021026 | -3.3257e-06 | 0.036
20211129 | 0118 | 0212 | 5M-9624 | 11/30/2020 | 0.00027441 | -795.63 | 1.09 | -0.0045878 | 0.00021026 | -3.3257e-06 | 0.036
20211215 | 0118 | 0212 | 5M-9624 | 11/30/2020 | 0.00027441 | -795.63 | 1.09 | -0.0045878 | 0.00021026 | -3.3257e-06 | 0.036
20220107 | 0118 | 0212 | 5M-9624 | 11/30/2020 | 0.00027441 | -795.63 | 1.09 | -0.0045878 | 0.00021026 | -3.3257e-06 | 0.036
20220124 | 0118 | 0212 | 5M-9624 | 11/30/2020 | 0.00027441 | -795.63 | 1.09 | -0.0045878 | 0.00021026 | -3.3257e-06 | 0.036
20220310 | 0282 | 3800 | 5M-9757 | 01/21/2021 | 0.00029409 | -832.63 | 1.46 | -0.0040976 | 0.00020421 | -3.4086e-06 | 0.036 | [10 empty] | Sensors swapped out in FEB2022, update cal numbers!
20220420 | 0282 | 3800 | 5M-9757 | 01/21/2021 | 0.00029409 | -832.63 | 1.46 | -0.0040976 | 0.00020421 | -3.4086e-06 | 0.036
20220516 | 0282 | 3800 | 5M-9757 | 01/21/2021 | 0.00029409 | -832.63 | 1.46 | -0.0040976 | 0.00020421 | -3.4086e-06 | 0.036
20220524 | 0282 | 3800 | 5M-9757 | 01/21/2021 | 0.00029409 | -832.63 | 1.46 | -0.0040976 | 0.00020421 | -3.4086e-06 | 0.036
20220602 | 0282 | 3800 | 5M-9757 | 01/21/2021 | 0.00029409 | -832.63 | 1.46 | -0.0040976 | 0.00020421 | -3.4086e-06 | 0.036 | [10 empty] | Poor pump performance swap sensors
20220616 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036 | [10 empty] | Freshly calibrated sensors installed before deployment
20220719 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036 | [10 empty] | working well phew!
20220909 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20220919 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20221028 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20221103 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20221209 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20221214 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230118 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230215 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230302 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230406 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230512 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230605 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230710 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230731 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20230908 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20231013 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20231114 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20231208 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20231220 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20240123 | 0164 | 3233 | 5M-8405 | 06/03/2022 | 0.00033395 | -828.96 | 1.37 | -0.0051884 | 0.00018386 | -2.6288e-06 | 0.036
20240215 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036 | [10 empty] | GPCTD, pump and DO swapped. Tubing replaced.
20240311 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20240418 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20240522 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20240618 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20240719 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20240813 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20240911 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20241024 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20241108 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20241216 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20250107 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20250131 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20250403 | 0317 | 3810 | 5M-10013 | 04/28/2023 | 0.00029398 | -837.12 | 1 | -0.0039114 | 0.00016467 | -2.6294e-06 | 0.036
20250514 | 0065 | 2422 | 5M-9822 | 12/02/2024 | 0.00022029 | -805.73 | 0.97 | -0.0033516 | 0.00020332 | -3.3366e-06 | 0.036 | [10 empty] | Freshly calibrated GPCTD installed.
20250606 | 0065 | 2422 | 5M-9822 | 12/02/2024 | 0.00022029 | -805.73 | 0.97 | -0.0033516 | 0.00020332 | -3.3366e-06 | 0.036
20250717 | 0065 | 2422 | 5M-9822 | 12/02/2024 | 0.00022029 | -805.73 | 0.97 | -0.0033516 | 0.00020332 | -3.3366e-06 | 0.036
20250813 | 0065 | 2422 | 5M-9822 | 12/02/2024 | 0.00022029 | -805.73 | 0.97 | -0.0033516 | 0.00020332 | -3.3366e-06 | 0.036
20250904 | 0065 | 2422 | 5M-9822 | 12/02/2024 | 0.00022029 | -805.73 | 0.97 | -0.0033516 | 0.00020332 | -3.3366e-06 | 0.036
20251020 | 0065 | 2422 | 5M-9822 | 12/02/2024 | 0.00022029 | -805.73 | 0.97 | -0.0033516 | 0.00020332 | -3.3366e-06 | 0.036
wgSparky
20160523 | 0173 | 3133 | N/A | 42089 | 0.00029344 | -868.83 | 2.48 | -0.0030498 | 0.00015794 | -2.5151e-06 | 0.036 | 0150 | 3009 | 2015-02-20 | 0.00027646 | -799.5 | 1.62 | -0.0032323 | 0.00015442 | -2.4218e-06 | 0.036 | *O2 is higher at the sub*
20160630 | 0173 | 3133 | N/A | 42089 | 0.00029344 | -868.83 | 2.48 | -0.0030498 | 0.00015794 | -2.5151e-06 | 0.036 | 0150 | 3009 | 2015-02-20 | 0.00027646 | -799.5 | 1.62 | -0.0032323 | 0.00015442 | -2.4218e-06 | 0.036 | *O2 is higher at the sub*
20180531 | 0173 | 3133 | N/A | 2/20/15 | 0.00029344 | -868.83 | 2.48 | -0.0030498 | 0.00015794 | -2.5151e-06 | 0.036 | 0150 | 3009 | 2015-02-20 | 0.00027646 | -799.5 | 1.62 | -0.0032323 | 0.00015442 | -2.4218e-06 | 0.036 | Glider upgraded to v200 and removed sub mounted GPCTD, waiting on new brackets
20180821 | 0173 | 3133 | N/A | 2/20/15 | 0.00029344 | -868.83 | 2.48 | -0.0030498 | 0.00015794 | -2.5151e-06 | 0.036 | 0150 | 3009 | 2015-02-20 | 0.00027646 | -799.5 | 1.62 | -0.0032323 | 0.00015442 | -2.4218e-06 | 0.036 | sub CTD remouted
20180905 | 0173 | 3133 | N/A | 2/20/15 | 0.00029344 | -868.83 | 2.48 | -0.0030498 | 0.00015794 | -2.5151e-06 | 0.036 | 0150 | 3009 | 2015-02-20 | 0.00027646 | -799.5 | 1.62 | -0.0032323 | 0.00015442 | -2.4218e-06 | 0.036 | Modified sub gpctd sampling on 17SEP2018
20181106 | 0173 | 3133 | N/A | 2/20/15 | 0.00029344 | -868.83 | 2.48 | -0.0030498 | 0.00015794 | -2.5151e-06 | 0.036 | 0150 | 3009 | 2015-02-20 | 0.00027646 | -799.5 | 1.62 | -0.0032323 | 0.00015442 | -2.4218e-06 | 0.036
20181113 | 0173 | 3133 | N/A | 2/20/15 | 0.00029344 | -868.83 | 2.48 | -0.0030498 | 0.00015794 | -2.5151e-06 | 0.036 | 0150 | 3009 | 2015-02-20 | 0.00027646 | -799.5 | 1.62 | -0.0032323 | 0.00015442 | -2.4218e-06 | 0.036 | Need to swap GPCTD this winter! Vehicle Sparky was damged and lost during this deployment
wgHansen
20190415 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | Shakedown mission
20190522 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036
20190625 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036
20190812 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036
20190927 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | Float mounted GPCTD stolen on this deployment!
20191209 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20191218 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20200619 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20200716 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20200909 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20201005 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20201113 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20201204 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20210218 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20210319 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20210409 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20210518 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20210625 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20210810 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20211026 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220118 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220202 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220209 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220216 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220323 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220420 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220525 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 0.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220705 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 1.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220815 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 2.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220908 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 3.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20220929 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 4.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20221020 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 5.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20221109 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 6.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20221130 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230112 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230209 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230404 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230515 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230607 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230707 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230809 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230901 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20230927 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20231018 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20231031 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20231108 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20231130 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20240118 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20240228 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20240415 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20240516 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20240708 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20240807 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20240911 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20241007 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20241119 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20250107 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20250311 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20250415 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20250527 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20250709 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20250820 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
20251001 | 0011 | 0244 | N/A | 8/9/2018 | 0.00029807 | -870.63 | 1.29 | -0.0038861 | 0.00018769 | -2.7361e-06 | 7.036 | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | No GPCTD+DO, copied values to get script to run
wgTest
20230509 | [11 empty] | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | Float mounted GPCTD installed, grab serial numbers.
20231129 | [11 empty] | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | Float mounted GPCTD installed, grab serial numbers.
20240304 | [11 empty] | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | Float mounted GPCTD installed, grab serial numbers.
20240321 | [11 empty] | 0206 | 3522 | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | Float mounted GPCTD installed, grab serial numbers.
20250604 | 0206 | 3522 | N/A | 2017-03-07 | 0.0003487 | -814.11 | 1.31 | -0.0043322 | 0.00021723 | -3.1952e-06 | 0.036 | [10 empty] | Float mounted GPCTD installed, grab serial numbers.

## MET
Deployment | Inventory info | Comments
S/N
wgOA
20120316 | N/A
20120521 | N/A
20120829 | N/A
20121015 | N/A
20130123 | N/A
20130315 | N/A
20130506 | N/A
20130625 | N/A
20130909 | N/A
20141217 | N/A
20150825 | N/A
20150923 | N/A
wgTiny
20140318 | N/A
20140403 | N/A
20140521 | N/A
20140619 | N/A
20140918 | N/A
20150212 | N/A
20150421 | N/A
20150501 | N/A
20150611 | N/A
20150824 | N/A
20150921 | N/A
20151019 | N/A
20151209 | N/A
20160128 | N/A | No MET data collected this deployment
20160225 | N/A | No science data collected this deployment
20160315 | N/A
20160627 | N/A | Switched to newer Airmar (200WX) at some point, not sure if it was for this deployment?
20160701 | N/A
20160830 | N/A
20160920 | N/A
20161005 | N/A
20161031 | N/A
20161212 | N/A | Air T low, atm press high second half of depl.
20170109 | 200WX 3589214 | Swapped in
20170307 | 200WX 3589214
20170412 | 200WX 3589214 | Air T jumped way up and is staying high.
20170605 | 200WX 3589214 | Air T still bad.
20170630 | 200WX 3589214 | Air T still bad, didn't have any spare Airmars at time of depl.
20170823 | 200WX 3240950 | Recently repaired by LR
20170921 | 200WX 3240950
20171019 | 200WX 3240950
20171212 | 200WX 3240950
20180116 | 200WX 3240950
20180208 | 200WX 3240950
20180323 | 200WX 3240950
20180516 | 200WX 3240950
20180710 | 200WX 3240950
20180730 | 200WX 3240950
20180829 | 200WX 3240950
20181001 | 200WX 3240950
20181030 | 200WX 3240950
20181210 | 200WX 3240950
20181217 | 200WX 3240950
20190114 | 200WX 3240950
20190213 | 200WX 3240950
20190304 | 200WX 3240950
20190502 | 200WX 3240950
20190513 | 200WX 3240950
20190627 | 200WX 3240950
20190719 | 200WX 3240950
20190909 | 200WX 3240950
20191001 | 200WX 3240950
20191114 | 200WX 3240950
20191213 | 200WX 3240950
20200106 | 200WX 3240950
20200203 | 200WX 3240950
20200227 | 200WX 3240950
20200708 | 200WX 3240950
20200717 | 200WX 3240950
20200903 | 200WX 3240950
20201006 | 200WX 3240950
20201202 | 200WX 3240950
20210209 | 200WX 3240950
20210318 | 200WX 3240950 | Airmar failed during deployment MAR28
20210408 | 200WX 60237814 | Airmar replaced!
20210528 | 200WX 60237814
20210701 | 200WX 60237814
20210804 | 200WX 60237814
20210818 | 200WX 60237814
20210902 | 200WX 60237814
20210928 | 200WX 60237814
20211028 | 200WX 60237814
20211129 | 200WX 60237814
20211215 | 200WX 3887904 | Light failed, entire light mast replaced with new Airmar
20220107 | 200WX 3887904
20220124 | 200WX 3887904
20220310 | 200WX 3887904
20220420 | 200WX 2971922 | Moisture and salt crystals inside light dome, entire assembly replaced.
20220516 | 200WX 2971922
20220524 | 200WX 2971922
20220602 | 200WX 2971922
20220616 | 200WX 2971922
20220719 | 200WX 2971922
20220909 | 200WX 2971922
20220919 | 200WX 3538425 | Light mast and Airmar replaced due to water in light lense
20221028 | 200WX 3538425
20221103 | 200WX 3538425
20221209 | 200WX 3538425
20221214 | 200WX 3538425
20230118 | 200WX 3538425
20230215 | 200WX 3538425
20230302 | 200WX 3538425
20230406 | 200WX 3538425
20230512 | 200WX 3538425
20230605 | 200WX 3538425
20230710 | 200WX 3538425
20230731 | 200WX 3538425
20230908 | 200WX 3538425
20231013 | 200WX 3538425
20231114 | 200WX 3538425
20231208 | 200WX 3538425
20231220 | 200WX 3538425
20240123 | 200WX 3538425
20240215 | 200WX 3538425
20240311 | 200WX 3538425 | Bad air temp readings… Relace!
20240418 | 200WX 3240950 | Refurbished Airmar installed on 03APR2024
20240522 | 200WX 3240950
20240618 | 200WX 3240950
20240719 | 200WX 3240950
20240813 | 200WX 3240950
20240911 | 200WX 3240950
20241024 | 200WX 3240950
20241108 | 200WX 3240950
20241216 | 200WX 3240950
20250107 | 200WX 3240950
20250131 | 200WX 3240950
20250403 | 200WX 3240950
20250514 | 200WX 3240950
20250606 | 200WX 3240950
20250717 | 200WX 3240950
20250813 | 200WX 3240950
20250904 | 200WX 3240950
20251020 | 200WX 3240950
wgSparky
20160523 | N/A | 200WX?
20160630 | N/A
20180531 | 200WX 3552050 | Not sure if this was the orginal shippe with Sparky from LR?
20180821 | 200WX 3552050
20180905 | 200WX 3552050
20181106 | 200WX 3552050 | Air temp bad?
20181113 | 200WX 3589214 | New Airmar installed, vehicle Sparky was damged and lost during this deployment
wgHansen
20190415 | 200WX 4160629
20190522 | 200WX 4160629
20190625 | 200WX 4160629
20190812 | 200WX 4160629
20190927 | 200WX 4160629
20191209 | 200WX 4160629 | bad air pressure?
20191218 | 200WX 4160629 | bad air pressure, needs to be replaced
20200619 | 200WX 4160629
20200716 | 200WX 4160629 | bad air pressure, needs to be replaced
20200909 | 200WX 60237811 | airmar replaced
20201005 | 200WX 60237811
20201113 | 200WX 60237811
20201204 | 200WX 60237811
20210218 | 200WX 60237811
20210319 | 200WX 60237811
20210409 | 200WX 60237811
20210518 | 200WX 60237811
20210625 | 200WX 60237811
20210810 | 200WX 60237811
20211026 | 200WX 60237811
20220118 | 200WX 60237811
20220202 | 200WX 60237811
20220209 | 200WX 60237811
20220216 | 200WX 60237811 | Airmar issues on 25FEB. self corrected appears fine now
20220323 | 200WX 60237811
20220420 | 200WX 60237811
20220525 | 200WX 60237811
20220705 | 200WX 60237811
20220815 | 200WX 60237811
20220908 | 200WX 60237811
20220929 | 200WX 60237811
20221020 | 200WX 60237811
20221109 | 200WX 60237811
20221130 | 200WX 60237811
20230112 | 200WX 60237811
20230209 | 200WX 60237811
20230404 | 200WX 60237811
20230515 | 200WX 60237811
20230607 | 200WX 60237811
20230707 | 200WX 60237811
20230809 | 200WX 60237811
20230901 | 200WX 60237811
20230927 | 200WX 60237811
20231018 | 200WX 2971922
20231031 | 200WX 2971922
20231108 | 200WX 2971922
20231130 | 200WX 2971922
20240118 | 200WX 2971922
20240228 | 200WX 2971922 | Air temp reading bad. Failed CCU S4 port and also Airmar cable? Sensor reaplaced
20240415 | 200WX 4160629 | "New" Airmar appear to be working!
20240516 | 200WX 4160629
20240708 | 200WX 4160629
20240807 | 200WX 4160629
20240911 | 200WX 4160629
20241007 | 200WX 4160629
20241119 | 200WX 4160629
20250107 | 200WX 60170992 | Light failed and mast replace.
20250311 | 200WX 60170992
20250415 | 200WX 60170992
20250527 | 200WX 60170992
20250709 | 200WX 60170992
20250820 | 200WX 60170992
20251001 | 200WX 60170992
wgTest
20230509 | 200WX 60170989
20231129 | 200WX 60170989
20240304 | 200WX 60170989 | Bad pressure readings! Replace!
20240321 | 200WX 60170989 | Bad pressure readings! Replace!
20250604 | 200WX 3887904 | Used Airmar installed on 04APR2024