Documentation for CALSH5

Program:  calsh5

Purpose:  Takes data that was created by calibrating a set of ATLAS

temperature pods in a bath of regulated temperature, and

calculates the temperature coefficients for each pod.

History: Written by Jim Holbrook, NOAA PMEL, in August 1975 in FORTRAN 77

for VAX/VMS.  Ported to HP-UX by Bob Herlien, MBARI, March 1991.

Machine:  hp850.mbari.org

Executable file:  /usr/local/bin/calsh5

Archive file:  /usr/local/atlas/calsh5.tar

How to create:

cd /usr/local/atlas

tar xf calsh5.tar

make

make install

Note - you must be superuser in order to run "make install". This is

because it moves the executable to /usr/local/bin

Syntax:  calsh5 filename

where "filename" is the root name of a file with extension .dat

Input:  filename.dat

Outputs: filename.coef, filename.out, filename.tbl

For example, the command line:

calsh5 cable47

will take input from file cable47.dat, and produce files cable47.coef,

cable47.out, and cable47.tbl.

Input file format

Example file: cable47.dat

10   10    1

5.988 7.986 10.985 13.983 16.979 19.975 22.976 25.972 28.968 31.967

4703

40df 45cd 4e01 5726 615c 6cca 798c 87b6 976b a8d5

4704

42d6 47e9 5054 59bc 6438 6ff4 7d0c 8b93 9bb0 ad8a

4705

3f39 4401 4bf5 54d1 5eb4 69c3 7617 83c5 92ed a3b7

4706

40e5 45d7 4e07 5730 6166 6cd3 7995 87be 9772 a8db

4707

424d 4757 4fb8 5910 637e 6f2a 7c32 8aa9 9ab3 ac7d

4708

4035 4518 4d30 563e 6057 6ba2 783d 863a 95bc a6ec

4709

43c0 48eb 517a 5b0d 65bb 71ae 7f03 8dd0 9e3b b06e

4710

4521 4a61 5318 5cd3 67ad 73d2 815e 9067 a115 b38f

4711

444e 497e 521a 5bb7 6670 7271 7fd6 8eb5 9f32 b177

4712

43db 4903 5192 5b23 65cf 71bf 7f14 8de0 9e4a b07b

Explanation of input file fields:

The first line contains 3 values.  The first value (10) is the number

of thermistor pods being calibrated.  The second value (10) is the number

of temperatures they were calibrated at.  The third field is the type

of thermistor pod.  It should always be "1" for our application.

The second line contains the temperatures that the pods were calibrated at.

The number of values on the second line should agree with the number of

temperatures declared on the first line.

These lines are followed by data on the thermistor pods.  Each thermistor

pod is represented by two lines.  The first line of the set contains a

numeric identifier for the pod.  This is normally composed of the cable

number in the upper two digits, and the pod number in the lower two

digits (e.g. 4703 is cable 47, pod 3), but can in fact contain any number.

The second line contains the output from that pod for each bath

temperature.  There should be one hexadecimal value for each bath temperature.

The total number of two-line sets should be equal to the number of pods

that were declared on line 1.  (Kinda sounds like IRS instructions, doesn't

it?)

Output file formats

Example file: cable47.coef

4703   20-Mar-91   .1043663E-02   .5460217E-03   .1982204E-05

4704   20-Mar-91   .1066641E-02   .5385619E-03   .2224851E-05

4705   20-Mar-91   .9919862E-03   .5614333E-03   .1685950E-05

4706   20-Mar-91   .1038554E-02   .5478049E-03   .1951845E-05

4707   20-Mar-91   .1065434E-02   .5392347E-03   .2169368E-05

4708   20-Mar-91   .1019605E-02   .5537753E-03   .1821134E-05

4709   20-Mar-91   .1078958E-02   .5359684E-03   .2270110E-05

4710   20-Mar-91   .1101158E-02   .5280912E-03   .2503968E-05

4711   20-Mar-91   .1097438E-02   .5288253E-03   .2458696E-05

4712   20-Mar-91   .1094275E-02   .5299755E-03   .2407642E-05

Expanation of output:

This file contains the A, B, and C coefficients for each temperature pod.

Each line contains: the pod number, the date the program was run, and

the A, B, and C coefficients.

Example file: cable47.out

Thermistor Calibration of -  4703

Steinhart-Hart Equation Fit

Old-Style Board

R = 7.68E+08/(2.525*N - 7680)

A =   .1043663E-02

B =   .5460217E-03

C =   .1982204E-05

R          T        T(R)     RESIDUAL

40DF  16607     22421.60547    5.9880    5.9880     .0000

45CD  17869     20513.24414    7.9860    7.9816     .0044

4E01  19969     17968.39063   10.9850   10.9920    -.0070

5726  22310     15785.33496   13.9830   13.9856    -.0026

615C  24924     13899.67188   16.9790   16.9748     .0042

6CCA  27850     12260.29199   19.9750   19.9725     .0025

798C  31116     10834.00586   22.9760   22.9756     .0004

87B6  34742      9594.77637   25.9720   25.9736    -.0016

976B  38763      8514.73535   28.9680   28.9690    -.0010

A8D5  43221      7570.00537   31.9670   31.9666     .0004

-6.97387E-03 4.43261E-03

Maximum Residual =    .0070

Standard Error   =    .0038

1

Thermistor Calibration of -  4704

Steinhart-Hart Equation Fit

Old-Style Board

R = 7.68E+08/(2.525*N - 7680)

A =   .1066641E-02

B =   .5385619E-03

C =   .2224851E-05

R          T        T(R)     RESIDUAL

42D6  17110     21619.94727    5.9880    5.9881    -.0001

47E9  18409     19792.42383    7.9860    7.9833     .0027

5054  20564     17358.24609   10.9850   10.9887    -.0037

59BC  22972     15261.01660   13.9830   13.9861    -.0031

6438  25656     13449.75684   16.9790   16.9748     .0042

6FF4  28660     11872.64746   19.9750   19.9729     .0021

7D0C  32012     10498.93066   22.9760   22.9763    -.0003

8B93  35731      9304.49121   25.9720   25.9735    -.0015

9BB0  39856      8261.93750   28.9680   28.9694    -.0014

AD8A  44426      7349.58789   31.9670   31.9659     .0011

-3.73901E-03 4.20751E-03

Maximum Residual =    .0042

Standard Error   =    .0029

And etc for the other pods.

Expanation of output: This file shows the calculations that were used to

derive A, B, and C.  It is meant to be self-expanatory.

Example file: cable47.tbl

ATLAS

TEMPERATURE CALIBRATIONS

------------------------

Date:20-Mar-91               Counts/Resist/Temp Conv:

Observers:___________                R =

Calibration Program:  calsh5         T =

Temperature Range:__________________

-------------------------------------------------------

Comments regarding calibration:

(procedure changes,new standard, software changes, etc)

------------------------------------------------------

A          B          C       Max      Std

Sensor     *E-02      *E-03      *E-05     Resid    Error Comments

4703    .1043663   .5460217   .1982204    .0070    .0038

4704    .1066641   .5385619   .2224851    .0042    .0029

4705    .0991986   .5614333   .1685950    .0071    .0038

4706    .1038554   .5478049   .1951845    .0039    .0028

4707    .1065434   .5392347   .2169368    .0064    .0036

4708    .1019605   .5537753   .1821134    .0031    .0023

4709    .1078958   .5359684   .2270110    .0034    .0021

4710    .1101158   .5280912   .2503968    .0042    .0029

4711    .1097438   .5288253   .2458696    .0047    .0032

4712    .1094275   .5299755   .2407642    .0034    .0027

Expanation of output:

This file contains the results in yet another form.  It is meant to be

printed, filled in by the person performing the calibration, and archived

as a permanent record of the calibration.